## Overview
Inter-device comparison between the High Stress Permeameter (HSP) and standard fixed-ring oedometers
This device comparison study is based on a set of measurements run on kaolinite in deionized water.
This spreadsheet contains the data and sample calculations described in Garcia and Waite (2025).
In tabs that include templates where a user would drop in data to make a calculation, the user-input spaces are shaded in green. The output is given in green text.
Consolidation |  | Tabulated and plotted consolidation data from all tests, including compression and recompression indices.
Permeability |  | Tabulated and plotted permeability data from all tests, including direct-flow measurements in the HSP and log time-based measurements in the fixed-ring oedometers.
Grain Size |  | Tabulated and plotted grain size distributions for the kaolin used in this study, as measured with a Horiba laser diffraction unit (LA-960).
HSP Specimen Height Example |  | Calibration data and correction factors for obtaining the specimen height for measurements in the HSP.
Direct Permeability Example |  | Example calculation for the HSP, including the correction for the porous stones that cap the specimen during testing.
Log Time Permeability Example |  | Example calculation for the oedometers, including the correction for the porous stones that cape the specimen during testing.
Material:
Kaolin used in this study comes from a single batch of "Peerless 2" air-floated kaolin (Vanderbilt Minerals). The measured grain size distribution for this material is given in the "Grain Size" tab.
https://www.vanderbiltminerals.com/resources/Kaolin_Clay_Ceramics.pdf
References:
Budhu, M. (2011), Soil Mechanics and Foundations, 3rd Edition, John Wiley & Sons, Inc., Hoboken, NJ.
ISBN: 978-1-118-13604-1
Collias, E. E., M. R. Rona, D. A. McMannus, and J. S. Creager (1963), Machine processing of technical data, University of Washington, Technical report #87.
https://apps.dtic.mil/sti/tr/pdf/AD0423746.pdf
Garcia, A.V., W.F. Waite (2025) , Design and functionality analysis of the High-Stress Permeameter (HSP), a device developed for measuring mechanical and hydraulic properties of gas hydrate-bearing pressure core, Geotechnical Testing Journal
doi: https://www.doi.org/10.1520/GTJ20250043
Jang, J., S. Dai, J. Yoneda, W. F. Waite, L. A. Stern, L.-G. Boze, T. S. Collett, and P. Kumar (2019), Pressure core analysis of geomechanical and fluid flow properties of seals associated with gas hydrate-bearing reservoirs in the Krishna-Godavari Basin, offshore India, Marine and Petroleum Geology, 108, 537-550.
doi: https://doi.org/10.1016/j.marpetgeo.2018.08.015
Lemmon, E. W., M. O. McLinden, and D. G. Friend (2016), Thermophysical Properties of Fluid Systems, in NIST Chemistry WebBook, NIST Standard Reference Database Number 69, edited by P. J. Linstrom and W. G. Mallard, National Institute of Standards and Technology, Gaithersburg MD, 20899 (http://webbook.nist.gov/chemistry/fluid/).
doi: https://www.doi.org/10.18434/T4D303
Santamarina, J. C., S. Dai, M. Terzariol, J. Jang, W. F. Waite, W. J. Winters, J. Nagao, J. Yoneda, Y. Konno, T. Fujii, K. Suzuki (2015), Hydro-bio-geomechanical properties of hydrate-bearing sediments from Nankai Trough, Marine and Petroleum Geology, 66, 434-450.
doi: https://doi.org/10.1016/j.marpetgeo.2015.02.033
Shepard, P. (1954), Nomenclature based on sand-silt-clay ratios, J Sediment Res, 24(3), 151-158.
doi: https://www.doi.org/10.1306/D4269774-2B26-11D7-8648000102C1865D

## Consolidation
Consolidation test results
These compression index data are measured on kaolin in deionized water, following the methods described for the Effective Stress Cell (ESC) in Santamarina et al. (2015) and Jang et al. (2019).
The compression indices, Cc (virgin compression) and Cr (recompression), are the slopes of the void ratio versus log of the vertical effective stress as indicated in the consolidation results plot. Slopes are calculated between the bounds shown in the Cc and Cr results boxes. Data points for compression are shown in blue, recompression points are in orange.
| High Stress Permeameter (HSP) | [10 empty] | GeoTac Automated Load Frame | [6 empty] | ELE International Dead-Weight Frame
| Test 1 |  |  |  | Test 2 |  |  |  | Test 3 |  |  | Frame 1 |  |  |  | Frame 2
| Void Ratio [-] | Stress [ MPa] | Log10(Stress) [-] |  | Void Ratio [-] | Stress [ MPa] | Log10(Stress) [-] |  | Void Ratio [-] | Stress [ MPa] | Log10(Stress) [-] | Void Ratio [-] | Stress [ MPa] | Log10(Stress) [-] |  | Void Ratio [-] | Stress [ MPa] | Log10(Stress) [-] | Void Ratio [-] | Stress [ MPa] | Log10(Stress) [-]
| 1.0911 | 0.34 | -0.46852 |  | 1.3507 | 0.19702 | -0.70549 |  | 1.5484 | 0.11292 | -0.94723 | 1.5694 | 0.027819 | -1.5557 |  | 1.6544 | 0.014046 | -1.8525 | 1.4187 | 0.034074 | -1.4676
| 1.0392 | 0.44 | -0.35655 |  | 1.0399 | 0.42801 | -0.36855 |  | 1.12 | 0.36055 | -0.44304 | 1.4667 | 0.048062 | -1.3182 |  | 1.5306 | 0.034131 | -1.4668 | 1.2323 | 0.10222 | -0.99045
| 0.99645 | 0.54 | -0.26761 |  | 0.87122 | 1.0958 | 0.03975 |  | 0.89462 | 1.0261 | 0.011195 | 1.2917 | 0.10479 | -0.97967 |  | 1.3066 | 0.10253 | -0.98913 | 1.0189 | 0.34074 | -0.46758
| 0.8523 | 1 | 0 |  | 0.6715 | 3.5284 | 0.54758 |  | 0.79031 | 2.0324 | 0.30801 | 1.0697 | 0.34131 | -0.46684 |  | 1.0796 | 0.34131 | -0.46684 | 0.84475 | 1.0222 | 0.0095456
| 0.64781 | 3.41 | 0.53275 |  | 0.57795 | 6.1482 | 0.78875 |  |  |  |  | 0.8854 | 1.0211 | 0.0090835 |  | 0.89442 | 1.0211 | 0.0090835 | 0.8746 | 0.34074 | -0.46758
| 0.55996 | 6 | 0.77815 |  | 0.49578 | 10.1782 | 1.0077 |  |  |  |  | 0.91434 | 0.34131 | -0.46684 |  | 0.92755 | 0.34131 | -0.46684 | 0.83125 | 1.0222 | 0.0095456
| 0.57531 | 3.41 | 0.53275 |  | 0.51778 | 5.1312 | 0.71022 |  |  |  |  | 0.87277 | 1.0211 | 0.0090835 |  | 0.88173 | 1.0211 | 0.0090835 | 0.64289 | 3.4074 | 0.53242
| 0.55339 | 4.5 | 0.65321 |  | 0.54241 | 2.598 | 0.41464 |  |  |  |  | 0.68427 | 3.4131 | 0.53316 |  | 0.69236 | 3.4131 | 0.53316 | 0.55661 | 5.997 | 0.77794
| 0.55334 | 6 | 0.77815 |  | 0.57867 | 1.0805 | 0.033618 |  |  |  |  | 0.58108 | 5.9976 | 0.77798 |  | 0.58475 | 5.9976 | 0.77798 | 0.57376 | 3.4074 | 0.53242
| 0.47075 | 10 | 1 |  | 0.60453 | 0.57741 | -0.23852 |  |  |  |  | 0.63288 | 2.0367 | 0.30892 |  | 0.63488 | 2.0367 | 0.30892 | 0.63254 | 1.0222 | 0.0095456
| 0.4909 | 5 | 0.69897 |  | 0.622 | 0.32751 | -0.48477 |  |  |  |  | 0.70154 | 0.59976 | -0.22202 |  | 0.70454 | 0.59976 | -0.22202 | 0.68858 | 0.34074 | -0.46758
| 0.51635 | 2.5 | 0.39794 | [8 empty] | 0.78004 | 0.14046 | -0.85245 |  | 0.78209 | 0.14046 | -0.85245 | 0.7466 | 0.10222 | -0.99045
| 0.55309 | 1 | 0 | [15 empty] | 0.79048 | 0.034074 | -1.4676
| 0.58212 | 0.5 | -0.30103
| Compression Index, Cc (blue data) | [6 empty] | Figure 1 | [6 empty] | Figure 2 |  |  |  | Figure 3
| Test | Compression Index,Cc [-] | Standard Deviation, ± [-] | Lower Stress Bound [MPa] | Upper Stress Bound [MPa] | Aspect Ratio (Diam/Hgt) [-]
| HSP 1 | 0.421 | 0.01 | 0.34 | 10 | 0.8
| HSP 2 | 0.39533 | 0.0035009 | 0.42801 | 10.1782 | 0.8
| HSP 3 | 0.444 | 0.04 | 0.36055 | 2.0324 | 0.8
| GeoTac 1 | 0.421 | 0.01 | 0.048062 | 5.9976 | 2.5
| GeoTac 2 | 0.416 | 0.01 | 0.034131 | 5.9976 | 2.5
| ELE | 0.386 | 0.004 | 0.034074 | 5.997 | 3.3
Recompression Index, Cr (orange data)
| Test | Recompression Index,Cr [-] | Standard Deviation, ± [-] | Lower Stress Bound [MPa] | Upper Stress Bound [MPa] | Aspect Ratio (Diam/Hgt) [-]
| HSP 1 | 0.088975 | 0.05 | 3.41 | 6 | 0.8 | [10 empty] | Note: error bars are taken from the standard deviations associated with each Cc or Cr measurement
| HSP 2 | -
| HSP 3 | -
| GeoTac 1 | 0.087345 | - | 0.34131 | 1.0211 | 2.5
| GeoTac 2 | 0.096269 | - | 0.34131 | 1.0211 | 2.5
| ELE | 0.090875 | - | 0.34074 | 1.0222 | 3.3
Example Cc Calculation (HSP 2)
|  |  |  | Slope of virgin compression curve ((-1)*Cc) = > | -0.39533 | 0.8906
|  |  |  | Uncertainty in Cc +> | 0.0035009 | 0.0022552
[5 empty] | 0.99976 | 0.0039338
[5 empty] | 12751.0848 | 3
[5 empty] | 0.19732 | 4.6423e-05
Note on B26: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: For an example of how the slope and slope uncertainty are calculated using the “LINEST” function, see the box for “Example Cc calculation (HSP2)” given below.

## Permeability
Permeability test results
Permeability results for the HSP are calculated as shown in the "Direct Permeability Example" tab. For the GeoTac oedometers, permeability results are calculated as shown in the "Log Time Permeability Example" tab.
Conversion of permeability (in milliDarcies) to hydraulic conductivity (m/s) assumes a water density of 1 kg/cubic meter, viscosity of 0.001 Pa·s, and gravity of 9.8 m/s^2.
| High Stress Permeameter (HSP) | [11 empty] | GeoTac Automated Load Frame
| Test 1 |  |  |  | Test 2 |  |  |  | Test 3 |  |  |  | Frame 1 | [5 empty] | Frame 2
| Void Ratio [-] | Stress [ MPa] | Permeability [mD] | Hydraulic Conductivity [m/s] | Void Ratio [-] | Stress [ MPa] | Permeability [mD] | Hydraulic Conductivity [m/s] | Void Ratio [-] | Stress [ MPa] | Permeability [mD] | Hydraulic Conductivity [m/s] | Void Ratio [-] | Stress [ MPa] | Mid-Step Void Ratio [-] | Mid-Step Stress [MPa] | Permeability [mD] | Hydraulic Conductivity [m/s] | Void Ratio [-] | Stress [ MPa] | Mid-Step Void Ratio [-] | Mid-Step Stress [MPa] | Permeability [mD] | Hydraulic Conductivity [m/s]
| 1.0911 | 0.34 |  |  | 1.3507 | 0.19702 |  |  | 1.5484 | 0.11292 |  |  | 1.6309 | 0.01389 | † |  |  |  | 1.7502 | 0.00072 | †
| 1.0392 | 0.44 |  |  | 1.0399 | 0.42801 |  |  | 1.12 | 0.36055 |  |  | 1.5694 | 0.027819 | 1.6002 | 0.020855 | 0.33596 | 3.2493e-09 | 1.6544 | 0.013787 | 1.7023 | 0.0072537 | 0.52941 | 5.1203e-09
| 0.99645 | 0.54 |  |  | 0.87122 | 1.0958 | 0.039564 | 3.8266e-10 | 0.89462 | 1.0261 | 0.034985 | 3.3837e-10 | 1.4667 | 0.048062 | 1.5181 | 0.037941 | 0.26046 | 2.5192e-09 | 1.5306 | 0.034081 | 1.5925 | 0.023934 | 0.26995 | 2.6109e-09
| 0.8523 | 1 | 0.045844 | 4.4339e-10 | 0.6715 | 3.5284 | 0.015993 | 1.5468e-10 | 0.79031 | 2.0324 | 0.025354 | 2.4522e-10 | 1.2917 | 0.10479 | 1.3792 | 0.076428 | 0.22722 | 2.1976e-09 | 1.3066 | 0.10295 | 1.4186 | 0.068517 | 0.21044 | 2.0354e-09
| 0.64781 | 3.41 | 0.017268 | 1.6701e-10 | 0.57795 | 6.1482 | 0.010971 | 1.0611e-10 |  |  |  |  | 1.0697 | 0.34131 | 1.1807 | 0.22305 | 0.11365 | 1.0992e-09 | 1.0796 | 0.34463 | 1.1931 | 0.22379 | 0.10702 | 1.0351e-09
| 0.55996 | 6 | 0.010005 | 9.6765e-11 | 0.49578 | 10.1782 | 0.0066293 | 6.4118e-11 |  |  |  |  | 0.8854 | 1.0211 | 0.97755 | 0.68123 | 0.051932 | 5.0228e-10 | 0.89442 | 1.0211 | 0.98702 | 0.68288 | 0.058573 | 5.6651e-10
| 0.57531 | 3.41 | 0.011169 | 1.0803e-10 | 0.51778 | 5.1312 | 0.0075154 | 7.2688e-11 |  |  |  |  | 0.68427 | 3.4131 | 0.78483 | 2.2171 | 0.023884 | 2.31e-10 | 0.69236 | 3.4131 | 0.79339 | 2.2171 | 0.027183 | 2.6291e-10
| 0.55339 | 4.5 | 0.011406 | 1.1031e-10 | 0.54241 | 2.598 | 0.0084469 | 8.1697e-11 |  |  |  |  | 0.58108 | 5.9976 | 0.63267 | 4.7054 | 0.013075 | 1.2646e-10
| 0.55334 | 6 | 0.010463 | 1.012e-10 | 0.57867 | 1.0805 | 0.010972 | 1.0612e-10
| 0.47075 | 10 | 0.0058449 | 5.6531e-11 | 0.60453 | 0.57741 | 0.012807 | 1.2387e-10 | [6 empty] | † | Note: For one dimensional consolidations, permeability is calculated from the rate of consolidation (fluid loss) during each consolidation step.
| 0.4909 | 5 | 0.0071711 | 6.9358e-11 | 0.622 | 0.32751 | [9 empty] | The void ratio and load are taken as the average value during each step as the load shifts from being supported by
| 0.51635 | 2.5 | 0.007828 | 7.5711e-11 | [11 empty] | the fluid (initially) to being supported by the compacted sediment grains (at the end of the step).
| 0.55309 | 1 | 0.012757 | 1.2338e-10
| 0.58212 | 0.5 | 0.012738 | 1.232e-10
| Figure 4 | [5 empty] | Figure 5
Note on U10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This is the test used as an example in the “Log Time Permeability Example” tab
Note on N11: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This is the test used as an example in the “HSP Specimen Height Example” and “Direct Permeability Example” tabs

## Grain Size
Grain Size Test Results
Measurement process summarized below the
data plots, and example calcs are to the right. |  |  |  | Classification |  |  |  | Method of Moments Statistics - Arithmetic (μm) † | [9 empty] | Grain Size Distribution: in bold are the minimum diameters for each "bin" (μm), and the data in each bin column represents the percentage of the grains that have diameters larger than the bin size, but smaller than the next largest bin size.
ANALYSIS_ID | SAMPLE_ID | ANALYSIS COMPLETION DATE | WEIGHT | SAND % (> 62.5µm) | SILT % (3.9<X<62.5µm) | CLAY % (< 3.9µm) | CLASSIFICATION Shepard (1954) | MEAN | STDEV | SKEWNESS | KURTOSIS | D10 | D25 | MEDIAN (D50) | D75 | D90 |  | 0.011 | 0.013 | 0.015 | 0.017 | 0.02 | 0.023 | 0.026 | 0.03 | 0.034 | 0.039 | 0.044 | 0.051 | 0.058 | 0.067 | 0.076 | 0.087 | 0.1 | 0.115 | 0.131 | 0.15 | 0.172 | 0.197 | 0.226 | 0.259 | 0.296 | 0.339 | 0.389 | 0.445 | 0.51 | 0.584 | 0.669 | 0.766 | 0.877 | 1.005 | 1.151 | 1.318 | 1.51 | 1.729 | 1.981 | 2.269 | 2.599 | 2.976 | 3.409 | 3.905 | 4.472 | 5.122 | 5.867 | 6.72 | 7.697 | 8.816 | 10.097 | 11.565 | 13.246 | 15.172 | 17.377 | 19.904 | 22.797 | 26.111 | 29.907 | 34.255 | 39.234 | 44.938 | 51.471 | 58.953 | 67.523 | 77.34 | 88.583 | 101.46 | 116.21 | 133.103 | 152.453 | 174.616 | 200
GS-005003 | LAB-Kaolin-Standard 1 | 2022-04-20 | 2.0439 | 0.42879 | 61.9483 | 37.6259 | Clayey silt | 11.3243 | 12.8979 | 1.5312 | 5.4425 | 0.22577 | 0.33278 | 7.8549 | 16.7668 | 29.6486 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.1795 | 0.37 | 0.6675 | 1.081 | 1.6865 | 2.514 | 3.532 | 4.6115 | 5.4835 | 5.7025 | 4.531 | 2.5165 | 1.1005 | 0.421 | 0.165 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.1445 | 0.225 | 0.348 | 0.526 | 0.7655 | 1.0675 | 1.431 | 1.8515 | 2.3245 | 2.838 | 3.371 | 3.8865 | 4.3485 | 4.714 | 4.808 | 4.6475 | 4.347 | 3.9845 | 3.6315 | 3.3265 | 3.063 | 2.787 | 2.418 | 1.911 | 1.3075 | 0.772 | 0.404 | 0.192 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
GS-005011 | LAB-Kaolin-Standard 1 (replicate) | 2022-04-20 | 2.0439 | 1.8786 | 65.2873 | 32.8386 | Clayey silt | 13.2145 | 15.6312 | 2.1292 | 9.0866 | 0.28522 | 0.42668 | 9.0909 | 18.1334 | 32.7258 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.1275 | 0.254 | 0.431 | 0.714 | 1.155 | 1.811 | 2.7305 | 3.9245 | 5.23 | 5.833 | 4.152 | 2.1035 | 0.8405 | 0.314 | 0.124 | 0 | 0 | 0 | 0 | 0 | 0 | 0.1465 | 0.225 | 0.348 | 0.5285 | 0.774 | 1.085 | 1.46 | 1.895 | 2.385 | 2.9215 | 3.4835 | 4.0355 | 4.531 | 4.95 | 5.174 | 5.132 | 4.8615 | 4.423 | 3.9065 | 3.395 | 2.9415 | 2.5575 | 2.2215 | 1.899 | 1.5505 | 1.1925 | 0.866 | 0.596 | 0.389 | 0.2435 | 0.1425 | 0 | 0 | 0 | 0 | 0
GS-005004 | LAB-Kaolin-Standard 2 | 2022-04-20 | 2.0998 | 0.26467 | 55.5662 | 44.1722 | Clayey silt | 9.128 | 10.9477 | 1.74 | 7.0804 | 0.22992 | 0.3187 | 6.1411 | 14.2934 | 23.2638 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.14 | 0.2945 | 0.5625 | 0.94 | 1.5135 | 2.3735 | 3.5755 | 5.0795 | 6.598 | 7.4355 | 6.1585 | 3.6025 | 1.7025 | 0.6825 | 0.272 | 0.119 | 0 | 0 | 0 | 0 | 0 | 0.1115 | 0.169 | 0.2505 | 0.3665 | 0.525 | 0.7295 | 0.982 | 1.2855 | 1.646 | 2.0675 | 2.553 | 3.091 | 3.65 | 4.197 | 4.6865 | 4.9645 | 4.9255 | 4.637 | 4.153 | 3.5525 | 2.914 | 2.2995 | 1.745 | 1.268 | 0.878 | 0.5795 | 0.3685 | 0.2255 | 0.1325 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
GS-005012 | LAB-Kaolin-Standard 2 (replicate) | 2022-04-20 | 2.0998 | 0.69432 | 55.4686 | 43.8401 | Clayey silt | 9.9359 | 12.6609 | 1.9951 | 8.1067 | 0.23118 | 0.32175 | 6.1195 | 14.69 | 25.994 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.1395 | 0.293 | 0.5565 | 0.927 | 1.488 | 2.3255 | 3.4945 | 4.96 | 6.45 | 7.2945 | 6.0785 | 3.579 | 1.7045 | 0.6915 | 0.2795 | 0.124 | 0 | 0 | 0 | 0 | 0 | 0.122 | 0.186 | 0.2765 | 0.406 | 0.583 | 0.8095 | 1.084 | 1.4055 | 1.774 | 2.1885 | 2.6435 | 3.1255 | 3.6025 | 4.046 | 4.421 | 4.59 | 4.4735 | 4.168 | 3.7405 | 3.266 | 2.805 | 2.3865 | 2.0095 | 1.6565 | 1.3115 | 0.978 | 0.6805 | 0.445 | 0.274 | 0.1595 | 0 | 0 | 0 | 0 | 0 | 0 | 0
GS-005159 | LAB-Kaolin-Standard 3 | 2023-06-30 | 4.4479 | 0.39488 | 68.5521 | 31.0561 | Clayey silt | 12.0772 | 12.1832 | 1.4472 | 5.6252 | 0.29326 | 0.44166 | 9.6609 | 17.6536 | 28.2758 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.1135 | 0.226 | 0.3865 | 0.6435 | 1.0485 | 1.656 | 2.5235 | 3.678 | 4.9865 | 5.8015 | 4.1895 | 2.0595 | 0.7875 | 0.278 | 0.103 | 0 | 0 | 0 | 0 | 0 | 0 | 0.111 | 0.1745 | 0.277 | 0.4325 | 0.652 | 0.9395 | 1.298 | 1.7315 | 2.243 | 2.833 | 3.4865 | 4.1675 | 4.8155 | 5.405 | 5.804 | 5.8975 | 5.6835 | 5.2115 | 4.5865 | 3.9125 | 3.2615 | 2.66 | 2.1015 | 1.5695 | 1.0695 | 0.6535 | 0.3615 | 0.183 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Average |  |  | 0.73225 | 61.3645 | 37.9066 | Clayey silt | 11.136 | 12.8642 | 1.7685 | 7.0683 | 0.25307 | 0.36832 | 7.7735 | 16.3074 | 27.9816 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0918 | 0.2397 | 0.4533 | 0.7531 | 1.2091 | 1.8833 | 2.8138 | 3.981 | 5.2268 | 6.1298 | 5.6805 | 3.6079 | 1.7341 | 0.6846 | 0.2617 | 0.094 | 0 | 0 | 0 | 0 | 0 | 0.0467 | 0.1514 | 0.2303 | 0.3491 | 0.519 | 0.7461 | 1.0316 | 1.376 | 1.7796 | 2.2417 | 2.7578 | 3.3115 | 3.8684 | 4.3876 | 4.8353 | 5.0681 | 5.0152 | 4.7394 | 4.3025 | 3.7886 | 3.2706 | 2.7904 | 2.3518 | 1.9331 | 1.5138 | 1.097 | 0.7334 | 0.4604 | 0.2755 | 0.1097 | 0.0487 | 0.0285 | 0 | 0 | 0 | 0 | 0
|  |  | Figure 6 | [6 empty] | Figure 7
† Method of Moments - definitions and example calculations from GS-005003 (LAB-Kaolin-Standard 1)
Step 1: Calculated quantities for each bin that are required to calculate the Method of Moments Statistics.
[18 empty] | Midpoint value for the bin (µm) | 0.012 | 0.014 | 0.016 | 0.0185 | 0.0215 | 0.0245 | 0.028 | 0.032 | 0.0365 | 0.0415 | 0.0475 | 0.0545 | 0.0625 | 0.0715 | 0.0815 | 0.0935 | 0.1075 | 0.123 | 0.1405 | 0.161 | 0.1845 | 0.2115 | 0.2425 | 0.2775 | 0.3175 | 0.364 | 0.417 | 0.4775 | 0.547 | 0.6265 | 0.7175 | 0.8215 | 0.941 | 1.078 | 1.2345 | 1.414 | 1.6195 | 1.855 | 2.125 | 2.434 | 2.7875 | 3.1925 | 3.657 | 4.1885 | 4.797 | 5.4945 | 6.2935 | 7.2085 | 8.2565 | 9.4565 | 10.831 | 12.4055 | 14.209 | 16.2745 | 18.6405 | 21.3505 | 24.454 | 28.009 | 32.081 | 36.7445 | 42.086 | 48.2045 | 55.212 | 63.238 | 72.4315 | 82.9615 | 95.0215 | 108.835 | 124.6565 | 142.778 | 163.5345 | 187.308
[18 empty] | Bin fraction * Midpoint value (µm) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.016783 | 0.039775 | 0.082102 | 0.15188 | 0.27153 | 0.46383 | 0.74702 | 1.1183 | 1.5217 | 1.8105 | 1.6493 | 1.0494 | 0.52549 | 0.23029 | 0.10337 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.26805 | 0.47813 | 0.84703 | 1.4662 | 2.4439 | 3.9038 | 5.9937 | 8.8816 | 12.772 | 17.861 | 24.2999 | 32.0889 | 41.1216 | 51.0573 | 59.6456 | 66.0363 | 70.7453 | 74.2731 | 77.5343 | 81.3462 | 85.7916 | 89.4097 | 88.8482 | 80.4263 | 63.0274 | 42.6237 | 25.5482 | 13.9068 | 0 | 0 | 0 | 0 | 0 | 0 | 0
[18 empty] | Bin fraction * (Midpoint value (µm) - Mean (µm))^2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.2264 | 0.46552 | 0.8375 | 1.3521 | 2.1017 | 3.1197 | 4.3618 | 5.6632 | 6.6916 | 6.9085 | 5.443 | 2.9938 | 1.2948 | 0.48899 | 0.18883 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.12957 | 0.19041 | 0.27505 | 0.38333 | 0.50619 | 0.62755 | 0.72865 | 0.78884 | 0.79001 | 0.71826 | 0.57103 | 0.36577 | 0.1517 | 0.01147 | 0.056208 | 0.38675 | 1.0652 | 2.1328 | 3.6506 | 5.7345 | 8.5268 | 12.0076 | 15.6248 | 18.0835 | 17.784 | 14.8697 | 10.8879 | 7.1695 | 0 | 0 | 0 | 0 | 0 | 0 | 0
[18 empty] | Bin fraction * (Midpoint value (µm) - Mean (µm))^3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -254.2689 | -522.1621 | -938.1088 | -1512.1348 | -2346.177 | -3475.3143 | -4847.1651 | -6275.8095 | -7392.0335 | -7604.0525 | -5965.6526 | -3265.4668 | -1404.4021 | -526.9975 | -202.0058 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -122.6924 | -175.1632 | -244.5256 | -327.2402 | -411.624 | -481.1609 | -519.951 | -514.895 | -460.5577 | -361.34 | -235.0241 | -112.209 | -28.3342 | -0.56578 | 6.0774 | 111.5666 | 527.3093 | 1560.399 | 3660.15 | 7529.3034 | 14226.7092 | 24923.7681 | 39718.6668 | 55627.9307 | 65587.7176 | 65259.9325 | 56523.3687 | 43810.5793 | 0 | 0 | 0 | 0 | 0 | 0 | 0
[18 empty] | Bin fraction * (Midpoint value (µm) - Mean (µm))^4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2855.6355 | 5856.9723 | 10508.011 | 16911.3698 | 26191.01 | 38714.2059 | 53865.4366 | 69547.0845 | 81658.1022 | 83696.0665 | 65385.1708 | 35617.3316 | 15233.2287 | 5679.5946 | 2161.0119 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1161.8072 | 1611.3735 | 2173.8986 | 2793.5748 | 3347.2318 | 3689.1913 | 3710.2511 | 3360.8594 | 2684.9462 | 1817.8191 | 967.3055 | 344.2314 | 52.9217 | 0.27908 | 6.5711 | 321.8395 | 2610.3019 | 11416.2363 | 36697.5013 | 98857.7113 | 237368.7842 | 517335.8961 | 1009657.5983 | 1711211.317 | 2418890.0317 | 2864110.2207 | 2934338.8354 | 2677143.0942 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Step 2: Calculate the Method of Moments Statistics (formulas are taken from Collias et al. (1963))
[18 empty] | Mean |  | 11.3243 | µm | The average grain size, weighted by the volume of grains in each grain size bin ( because this is a laser-based particle distribution, grain volume rather than mass is tracked here).
[18 empty] | Standard Deviation |  | 12.8979 | µm | Indicates the width of the grain size distribution, which is associated with how well sorted (uniform) the grain sizes are. The high result for this specimen reflects it's bimodal grain size distribution.
[18 empty] | Skewness |  | 1.5312 |  | Indicates whether the greater volume of sediment lies above the mean grain size (positive skewness) or below the mean (negative skewness). The positive result here indicates slightly more of the specimen is in the larger grain size cluster.
[18 empty] | Kurtosis |  | 5.4425 |  | Like Standard Deviation, this indicates the spread of the grain size and the degree of sorting (low Kurtosis implies a well-sorted sediment). The high result here indicates the widely-spread, bimodal grain size distribution.
Step 3: The "percent finer" dxx values are taken from the cumulative grain size distribution
[18 empty] | Cumulative | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.1795 | 0.5495 | 1.217 | 2.298 | 3.9845 | 6.4985 | 10.0305 | 14.642 | 20.1255 | 25.828 | 30.359 | 32.8755 | 33.976 | 34.397 | 34.562 | 34.562 | 34.562 | 34.562 | 34.562 | 34.562 | 34.562 | 34.562 | 34.7065 | 34.9315 | 35.2795 | 35.8055 | 36.571 | 37.6385 | 39.0695 | 40.921 | 43.2455 | 46.0835 | 49.4545 | 53.341 | 57.6895 | 62.4035 | 67.2115 | 71.859 | 76.206 | 80.1905 | 83.822 | 87.1485 | 90.2115 | 92.9985 | 95.4165 | 97.3275 | 98.635 | 99.407 | 99.811 | 100.003 | 100.003 | 100.003 | 100.003 | 100.003 | 100.003 | 100.003 | 100.003
Figure 8
[18 empty] | d10 | 0.22577 | µm | 10 | 10 | 0.01 | 0.22577
[18 empty] | d25 | 0.33278 | µm | 25 | 25 | 0.01 | 0.33278
[18 empty] | Median (d50) | 7.8549 | µm | 50 | 50 | 0.01 | 7.8549
[18 empty] | d75 | 16.7668 | µm | 75 | 75 | 0.01 | 16.7668
[18 empty] | d90 | 29.6486 | µm | 90 | 90 | 0.01 | 29.6486
Ternary plot is based on the classification system defined by Shepard (1954)
Grain Size Measurement Process Steps
About 2-5 grams of dry kaolin are placed in a pre-weighed beaker and the gross weight is recorded. Because the kaolin is known from previous measurements to contain only particles smaller than 4 mm-diameter, samples are considered to be “Fine-fraction” only and are not pre-sieved.
Fine fractions ready for analysis by the Horiba laser diffraction unit (LA-960) are rehydrated with distilled water. Fifteen (15) ml of pre-mixed 40 g/l sodium hexametaphosphate [(NaPO3)6] are added to each sample. If the height of the fluid in the laser diffraction vial is less than 5 cm, more distilled water is added to raise the level to no more than 8 cm in the vial. The samples are gently stirred, covered, and allowed to soak for at least 1 hour (for samples that were not dried) up to 24 hours (for samples that were dried). Soaked vials are placed into an ultrasonic bath and run for 10 minutes at a frequency of 37 Hz with a power level of 100. If the samples appear to be fully disaggregated, they are placed into pre-determined autosampler locations, and are run using the Horiba LA-960 for Windows software to get the fine fraction grain size distributions.
Standard formulae are applied to the distribution to acquire the distribution statistics (mean, standard deviation, skewness, kurtosis) and the distribution markers (d10 through d90 and the median grain size, d50).
Note on T17: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The midpoint value (halfway between a given bin size and the bin size for the previous bin) is used in all of the Method of Moments calculations.
Note on T18: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This result is used in the summation for the Mean grain size.
Note on T19: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This result is used in the summation for the Standard Deviation of the Mean grain size.
Note on T20: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This result is used in the summation for the Skewness of the grain size distribution.
Note on T21: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This result is used in the summation for the Kurtosis of the grain size distribution.
Note on T32: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This is the cumulative fraction of grains that are smaller than the stated bin size (for GS-005003 (LAB-Kaolin-Standard 1). This is used to calculate the dxx values, including the Median (d50).

## HSP Specimen Height Example
Specimen Height Corrections For the High-Stress Permeameter | [9 empty] | Points for the elastic deformation fit curve
Specimen height is calculated from the piston protrusion (see diagram), with corrections for the chamber pressure and load.
These calibration data come from piston protrusion measurements made in the absence of a specimen.
[10 empty] | Effective Stress [MPa] | Elastic Deformation [mm]
| Initial Loading at 3 MPa chamber pressure † |  |  |  | Unloading at 3 MPa chamber pressure |  |  |  | Figure 10 | 0.1 | -0.047585
| Effective Stress [MPa] | LVDT [mm] | Deformation [mm] |  | Effective Stress [MPa] | LVDT [mm] | Deformation [mm] | Elastic Deformation, ∆d [mm] |  | 0.2 | -0.06291
| 0.01 | 6.191 | 0 |  | 16.037 | 5.752 | -0.439 | -0.3647 |  | 0.3 | -0.074071
| 0.014581 | 6.1661 | -0.0249 |  | 15.0854 | 5.762 | -0.429 | -0.3547 |  | 0.4 | -0.083172
| 0.21458 | 6.1173 | -0.0737 |  | 10.0854 | 5.8207 | -0.3703 | -0.296 |  | 0.5 | -0.090994
| 0.91458 | 6.049 | -0.142 |  | 5.0854 | 5.892 | -0.299 | -0.2247 |  | 0.6 | -0.097928
| 2.9146 | 5.963 | -0.228 |  | 4.9146 | 5.894 | -0.297 | -0.2227 |  | 0.7 | -0.1042
| 9.9146 | 5.8325 | -0.3585 |  | 1.0854 | 5.997 | -0.194 | -0.1197 |  | 0.8 | -0.10996
| 14.9146 | 5.7705 | -0.4205 |  | 1.0854 | 5.999 | -0.192 | -0.1177 |  | 0.9 | -0.1153
| 16.037 | 5.752 | -0.439 |  | 0.58542 | 6.021 | -0.17 | -0.0957 |  | 1 | -0.1203
| † Note: Load is measured with a NIST-traceable load cell, |  |  |  | 0.58542 | 6.021 | -0.17 | -0.0957 |  | 1.1 | -0.12501
| LVDT (Linear Voltage Displacement Transducer) is |  |  |  | 0.18542 | 6.056 | -0.135 | -0.0607 |  | 1.2 | -0.12947
| calibrated with polished gauge blocks. | [8 empty] | 1.3 | -0.13371
[10 empty] | 1.4 | -0.13776
| Reloading at 10 MPa chamber pressure |  |  |  | Unloading at 10 MPa chamber pressure |  |  |  |  | 1.5 | -0.14164
| Effective Stress [MPa] | LVDT [mm] | Deformation [mm] |  | Effective Stress [MPa] | LVDT [mm] | Deformation [mm] | Elastic Deformation, ∆d [mm] |  | 1.6 | -0.14537
| 0.21458 | 5.986 | -0.205 |  | 13.7146 | 5.6797 | -0.5113 | -0.354 |  | 1.7 | -0.14897
| 0.91458 | 5.932 | -0.259 |  | 10.0854 | 5.7186 | -0.4724 | -0.3151 |  | 1.8 | -0.15244
| 2.9146 | 5.851 | -0.34 |  | 5.0854 | 5.789 | -0.402 | -0.2447 |  | 1.9 | -0.15579
| 9.9146 | 5.7247 | -0.4663 |  | 1.0854 | 5.902 | -0.289 | -0.1317 |  | 2 | -0.15905
| 13.7146 | 5.6797 | -0.5113 |  | 0.38542 | 5.947 | -0.244 | -0.0867 | Figure 11 | 2.1 | -0.1622
[5 empty] | 0.18542 | 5.973 | -0.218 | -0.0607 |  | 2.2 | -0.16527
[10 empty] | 2.3 | -0.16826
| Corrections applied to the piston protrusion calculation for specimen length | [5 empty] | Figure 9 |  |  | 2.4 | -0.17116
| Original piston protrusion for a 0 mm specimen= |  |  | 77.5 | mm |  |  |  |  | 2.5 | -0.174
[10 empty] | 2.6 | -0.17677
| Permanent deformation d*(3 MPa)= |  |  | -0.0613 | mm |  |  |  |  | 2.7 | -0.17948
| Permanent deformation d*(10 MPa)= |  |  | -0.013 | mm |  |  |  |  | 2.8 | -0.18213
| Total permanent deformation d*(total)= |  |  | -0.0743 | mm |  |  |  |  | 2.9 | -0.18472
[10 empty] | 3 | -0.18726
| dp (10MPa-3Mpa) |  |  | -0.083 | mm |  |  |  |  | 3.1 | -0.18975
[10 empty] | 3.2 | -0.19219
| Elastic deformation, ∆d [mm] = 0.12s'^(.4) | [8 empty] | 3.3 | -0.19459
[10 empty] | 3.4 | -0.19695
| Example correction (User can change green cells): | [8 empty] | 3.5 | -0.19926
| Measured Piston Extrusion [mm] | Chamber Pressure [MPa] | Effective Stress [MPa] | Corrected specimen length [mm] | [5 empty] | 3.6 | -0.20153
| 118.52 | 10 | 2.03 | 41.2536 | [5 empty] | 3.7 | -0.20377
[10 empty] | 3.8 | -0.20597
[10 empty] | 3.9 | -0.20814
[10 empty] | 4 | -0.21027
[10 empty] | 4.1 | -0.21237
[10 empty] | 4.2 | -0.21444
[10 empty] | 4.3 | -0.21648
[10 empty] | 4.4 | -0.2185
[10 empty] | 4.5 | -0.22049
[10 empty] | 4.6 | -0.22245
[10 empty] | 4.7 | -0.22438
[10 empty] | 4.8 | -0.22629
[10 empty] | 4.9 | -0.22818
[10 empty] | 5 | -0.23004
[10 empty] | 5.1 | -0.23189
[10 empty] | 5.2 | -0.23371
[10 empty] | 5.3 | -0.23551
[10 empty] | 5.4 | -0.23729
[10 empty] | 5.5 | -0.23905
[10 empty] | 5.6 | -0.24079
[10 empty] | 5.7 | -0.24251
[10 empty] | 5.8 | -0.24422
[10 empty] | 5.9 | -0.2459
[10 empty] | 6 | -0.24757
[10 empty] | 6.1 | -0.24923
[10 empty] | 6.2 | -0.25087
[10 empty] | 6.3 | -0.25249
[10 empty] | 6.4 | -0.25409
[10 empty] | 6.5 | -0.25569
[10 empty] | 6.6 | -0.25726
[10 empty] | 6.7 | -0.25883
[10 empty] | 6.8 | -0.26038
[10 empty] | 6.9 | -0.26191
[10 empty] | 7 | -0.26343
[10 empty] | 7.1 | -0.26494
[10 empty] | 7.2 | -0.26644
[10 empty] | 7.3 | -0.26792
[10 empty] | 7.4 | -0.2694
[10 empty] | 7.5 | -0.27086
[10 empty] | 7.6 | -0.27231
[10 empty] | 7.7 | -0.27374
[10 empty] | 7.8 | -0.27517
[10 empty] | 7.9 | -0.27659
[10 empty] | 8 | -0.27799
[10 empty] | 8.1 | -0.27939
[10 empty] | 8.2 | -0.28077
[10 empty] | 8.3 | -0.28214
[10 empty] | 8.4 | -0.28351
[10 empty] | 8.5 | -0.28486
[10 empty] | 8.6 | -0.28621
[10 empty] | 8.7 | -0.28754
[10 empty] | 8.8 | -0.28887
[10 empty] | 8.9 | -0.29019
[10 empty] | 9 | -0.2915
[10 empty] | 9.1 | -0.2928
[10 empty] | 9.2 | -0.29409
[10 empty] | 9.3 | -0.29537
[10 empty] | 9.4 | -0.29665
[10 empty] | 9.5 | -0.29792
[10 empty] | 9.6 | -0.29917
[10 empty] | 9.7 | -0.30043
[10 empty] | 9.8 | -0.30167
[10 empty] | 9.9 | -0.30291
[10 empty] | 10 | -0.30413
[10 empty] | 10.1 | -0.30536
[10 empty] | 10.2 | -0.30657
[10 empty] | 10.3 | -0.30778
[10 empty] | 10.4 | -0.30898
[10 empty] | 10.5 | -0.31017
[10 empty] | 10.6 | -0.31136
[10 empty] | 10.7 | -0.31254
[10 empty] | 10.8 | -0.31371
[10 empty] | 10.9 | -0.31488
[10 empty] | 11 | -0.31604
[10 empty] | 11.1 | -0.31719
[10 empty] | 11.2 | -0.31834
[10 empty] | 11.3 | -0.31948
[10 empty] | 11.4 | -0.32062
[10 empty] | 11.5 | -0.32175
[10 empty] | 11.6 | -0.32287
[10 empty] | 11.7 | -0.32399
[10 empty] | 11.8 | -0.3251
[10 empty] | 11.9 | -0.32621
[10 empty] | 12 | -0.32731
[10 empty] | 12.1 | -0.32841
[10 empty] | 12.2 | -0.3295
[10 empty] | 12.3 | -0.33058
[10 empty] | 12.4 | -0.33166
[10 empty] | 12.5 | -0.33274
[10 empty] | 12.6 | -0.33381
[10 empty] | 12.7 | -0.33487
[10 empty] | 12.8 | -0.33593
[10 empty] | 12.9 | -0.33699
[10 empty] | 13 | -0.33804
[10 empty] | 13.1 | -0.33908
[10 empty] | 13.2 | -0.34012
[10 empty] | 13.3 | -0.34116
[10 empty] | 13.4 | -0.34219
[10 empty] | 13.5 | -0.34321
[10 empty] | 13.6 | -0.34424
[10 empty] | 13.7 | -0.34525
[10 empty] | 13.8 | -0.34627
[10 empty] | 13.9 | -0.34727
[10 empty] | 14 | -0.34828
[10 empty] | 14.1 | -0.34928
[10 empty] | 14.2 | -0.35027
[10 empty] | 14.3 | -0.35127
[10 empty] | 14.4 | -0.35225
[10 empty] | 14.5 | -0.35324
[10 empty] | 14.6 | -0.35422
[10 empty] | 14.7 | -0.35519
[10 empty] | 14.8 | -0.35616
[10 empty] | 14.9 | -0.35713
[10 empty] | 15 | -0.35809
[10 empty] | 15.1 | -0.35905
[10 empty] | 15.2 | -0.36001
[10 empty] | 15.3 | -0.36096
[10 empty] | 15.4 | -0.36191
[10 empty] | 15.5 | -0.36285
[10 empty] | 15.6 | -0.36379
[10 empty] | 15.7 | -0.36473
[10 empty] | 15.8 | -0.36567
[10 empty] | 15.9 | -0.3666
[10 empty] | 16 | -0.36752
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Deformation includes elastic deformation of the system and permanent deformation of the porous metal stones.
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This is the total deformation minus the permanent deformations measured after the two loading cycles.
Note on D20: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Deformation includes elastic deformation of the system and permanent deformation of the porous metal stones.
Note on I20: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This is the total deformation minus the permanent deformations measured after the two loading cycles. The additional change in piston protrusion due to the chamber pressure difference between 10 and 3 MPa is also removed to put this data on the same footing as the 3 MPa elastic deformation.
Note on B35: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This is the change in piston protrusion (a decrease) when the chamber swells as the chamber water pressure increases from 3 to 10 MPa.
Note on B37: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This is the power law fit result from plotting the elastic deformations at 3 and 10 MPa (Column I) as a function of effective stress (Column F).
Note on B41: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: To illustrate the test results for the HSP Test 3, 2.03 MPa stress datapoint in the “Permeability” tab, use a piston protrusion of 118.52 mm, chamber pressure of 10 MPa, and a vertical effective stress of 2.03 MPa. The resulting specimen height (41.25 mm) can then be used in the “Direct Permeability Process” tab to illustrate how permeability is calculated for the HSP.
Piston Protrusion

## Direct Permeability Example
High Stress Permeameter Direct Permeability Measurement Process Steps
These data show the permeability correction to account for the porous stones that cap the specimen (step 1) and how a measured permeability calculation is made in the HSP (step 2), and how the specimen permeability is calculated (step 3):
| Step 1: Porous stone permeability | [5 empty] | Step 2: Apparent Permeability = Permeability of the specimen and porous stones
Porous metal stone permeability - dependence on load |  |  |  |  | Calculating permeability from a series of flow rate versus pressure difference measurements
Note: Porous stone permeability is measured via direct flow (see Step 2), but with no specimen in place.
From Darcy's Law, permeability, k, can be calculated from the slope of measured flow rate, Q, versus pressure difference across the specimen, ΔP.
| Initial Loading† |  | Unloading†† |  |  | The example shown here is from HSP Test 3, when permeability was measured at ~2MPa vertical effective stress, s'.
| Effective Stress [MPa] | Permeability [mD] | Effective Stress [MPa] | Permeability [mD] |  | A = cross sectional area for flow (top plunger cap area) = |  |  | 0.002341 | m^2
| 0.01 | 1677.6494 | 16.037 | 1439.0768 |  | H = Specimen height = |  |  | 41.25 | mm | (ΔP/H is the pressure gradient)
| 0.014581 | 1662.1011 | 15.0854 | 1439.225 |  | µ = viscosity of water = |  |  | 0.00153 | Pa·s | (viscosity for water at 4.4 °C at 10 MPa pressure, as given by : |  |  | NIST Webbook)
| 0.91458 | 1531.9534 | 5.0854 | 1442.3256
| 9.9146 | 1455.7544 | 0.58542 | 1452.3733
| 16.037 | 1439.0768
[6 empty] | Measured Data |  |  | Figure 13
| † Porous stones are deforming during their initial loading, reducing their permeability. |  | †† Porous stones, deformed by the peak load, now have a repeatable stress-dependent permeability that depends on the elastic deformation in the stones. |  |  | Differential Pressure, ΔP [kPa] | Flow Rate, Q [µl/min]
[6 empty] | 11.4381 | 2.0594
| Figure 12 |  |  |  |  | 21.6683 | 3.0922
[6 empty] | 31.959 | 3.7284
[6 empty] | 43.9172 | 4.9552
[6 empty] | 64.0172 | 6.5067
[6 empty] | 106.131 | 10.0361
[6 empty] | 159.841 | 14.3525
[6 empty] | 213.6621 | 18.8544
[6 empty] | 266.7741 | 23.4316
Apparent permeability calculation example:
[6 empty] | Slope = | 0.083038 | µL/min·(1/kPa)
[7 empty] | 1.384e-15 | (m^3/s)·(1/Pa)
[6 empty] | km = | 3.7311e-17 | m^2
[7 empty] | 0.037806 | milliDarcy, mD
Step 3: Specimen Permeability: correcting for the porous stone effect
| Permeability correction to account for the HSP porous stones: |  |  |  |  | The measured specimen permeability is taken from the series equation for flow through the two porous stones and the specimen in series:
[11 empty] | Specimen height, Hs = | 41.25 | mm
[11 empty] | Total porous stone height, Hps = | 20.26 | mm
[11 empty] | Vertical effective stress, s' = | 2.03 | MPa
[11 empty] | Porous stone permeability, kps = | 1446.8235 | mD
[11 empty] | Measured apparent permeability, km = | 0.037806 | mD
[11 empty] | Specimen permeability, ks = | 0.025354 | mD
𝒌_((𝒆𝒍𝒂𝒔𝒕𝒊𝒄 𝒅𝒆𝒇𝒐𝒓𝒎𝒂𝒕𝒊𝒐𝒏))=𝟏𝟒𝟒𝟗.𝟗〖𝝈′〗^(−.𝟎𝟎𝟑)
𝑘_𝑠=𝐻_𝑠/((𝐻_𝑠+𝐻_𝑝𝑠)/𝑘_𝑚 −𝐻_𝑝𝑠/𝑘_𝑝𝑠 )
Note on N10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: To illustrate the HSP permeability calculation for HSP Test 3, 2.03 MPa vertical effective stress case from the “Permeability” tab, the specimen height here should be 41.25 mm, as calculated in the “HSP Specimen Height Example” tab.
Note on K16: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: These example data were measured for HSP Test 3, 2.03 MPa vertical effective stress (See “Permeability” tab for the final result, 0.0253 mD)
Note on Q46: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The data point for HSP Test 3, 2.03 MPa vertical effective stresses (.0253 mD) will be replicated here if the user has input the specimen height 41.25 mm in Cell N10.

## Log Time Permeability Example
Oedometer Log Time Permeability Measurement Process Steps
Terzaghi's log time permeability method used here follows the process given by Budhu (2011), and the porous stone correction process from the "Direct Permeability Example" tab.
Here, raw data for stress and specimen height are pasted into the green region (Columns B and C), then the row with the first loading point (row 6 in this example) is input by the user into cell G8. Follow step 1 to continue the permeability calculation.
Time | Stress | Height of Specimen | Step Time | Height of Specimen at the Start of the Load Step
[s] | [kPa] | [mm] | [s] | 24.7121 | [mm] |  | d0 [mm] | d100 [mm] | Δd [mm] | Step 1: Identify key data points in the compaction curve
0 | 27.819 | 24.7121 | 0 |  |  |  | 24.6603 | 23.7219 | 0.1786 |  | d1: Choose a point before the steepest part of the compaction curve (cell H8)
10 | 29.5773 | 24.711 | 10 | Row for Zero Time | Row # for t1, d1 | Row # for d100 | d1 [mm] | d2 [mm] | d50 [mm] |  | d100: Choose the final data point (100% consolidation, cell I8)
20 | 47.5432 | 24.5619 | 20 | 6 | 10 | 1887 | 24.4817 | 24.3031 | 24.1911 |  | d2: Spreadsheet chooses the point at time t2 = 4·t1
30 | 45.4404 | 24.5144 | 30 | Dynamic viscosity, µ |  |  | t1 [s] | t2 [s] | t50 [s] |  | d0: Spreadsheet calculates the change in height, Δd, between d1 and d2, then adds that height change to d1 to get the theoretical initial height of the specimen.
40 | 47.3362 | 24.4817 | 40 | 0.0009544 | [Pa-s] |  | 40 | 160 | 300 |  | d50: Spreadsheet chooses the point at the specimen height representing 50% compaction (halfway between d0 and d100).
50 | 47.355 | 24.4593 | 50 | Unit weight of water, γw
60 | 47.4162 | 24.4368 | 60 | 9.81 | [kN/m3] |  | σ0 [kPa] | σf [kPa] | Δσ [kPa] | Step 2: Calculate the modulus of volume compressibility, mv
70 | 47.7831 | 24.4185 | 70 | Thickness of top porous stone |  |  | 27.819 | 48.0693 | 20.2503 |  | mv = - (ΔH/H0)/Δσ
80 | 47.976 | 24.4026 | 80 | 6.4 | [mm] |  | ΔH[mm] | ΔH/H0[] | mv [1/kPa] |  | ΔH/H0, the strain during the load step, is calculated from the total height change, ΔH, and the initial height, H0
90 | 48.4135 | 24.3835 | 90 | Thickness of bottom porous stone |  |  | -0.93836 | -0.038051 | 0.0018791 |  |  | ΔH = d100-d0; H0 = d0
100 | 48.1313 | 24.3715 | 100 | 5.5 | [mm] | [5 empty] | Δσ, the change in stress is taken as the difference in the initial and final stresses, σf - σ0
110 | 48.7099 | 24.3584 | 110 | Permeability of porous stones
120 | 47.8443 | 24.3461 | 120 | 2955 | [mD] |  | Hdr [m] | Cv [m2/s] |  | Step 3: Calculate the coefficient of consolidation, Cv from the step 1 campaction curve points
130 | 48.4841 | 24.3334 | 130 | Permeability unit conversion |  |  | 0.018046 | 2.1384e-07 |  |  | Cv = 0.197·Hdr/t50, as per the definition for the log time method (Budhu, 2011).
140 | 48.1642 | 24.3312 | 140 | 1013249970000000 | [mDarcy/m2] | [5 empty] | Hdr, the drainage height, is the path water takes to exit the system, so it includes 2 elements:
150 | 48.4135 | 24.313 | 150 | [8 empty] | 1. Path through specimen = half of specimen height at 50% consolidation = d50/2
160 | 48.2018 | 24.3031 | 160 |  | Figure 14 | [6 empty] | 2. Path through either the top or bottom porous stone = average thickness of the porous stones
170 | 48.5829 | 24.2906 | 170 | [8 empty] | Hdr = d50/2 + Average porous stone thickness
180 | 47.6232 | 24.2818 | 180 | [7 empty] | t50 is the time associated with d50, the 50% consolidation point.
190 | 48.1642 | 24.2724 | 190
200 | 48.0748 | 24.2644 | 200 | [6 empty] | Step 4: Calculate the measured hydraulic conductivity, K [m/s], and convert to effective permeability, k [milliDarcy]
210 | 48.6958 | 24.2559 | 210 | [7 empty] | Km = mv·Cv·γw, where γw is the unit weight of water [kN/m3]
220 | 48.3194 | 24.2472 | 220 | [7 empty] | km = (K·µ/γw)·(1.01325e15 mDarcy/m2)
230 | 48.1501 | 24.2417 | 230 | [7 empty] | Km [m/s] | km [mDarcy]
240 | 48.1077 | 24.2365 | 240 | [7 empty] | 3.9418e-09 | 0.38857
250 | 48.0136 | 24.2228 | 250
260 | 47.9807 | 24.2143 | 260 | [6 empty] | Step 5: Calculate the specimen permeability from the measured permeability,
270 | 47.562 | 24.2037 | 270 | [6 empty] | by removing contributionsfrom the porous stones.
280 | 47.7502 | 24.1998 | 280 | [7 empty] | This porous stone correction uses the same parallel model presented in the "Direct Permeability Process" tab
290 | 47.435 | 24.1921 | 290 | [7 empty] | Specimen Permeability, ks= |  | 0.26046 | mD
300 | 47.6796 | 24.1865 | 300
310 | 48.2065 | 24.1782 | 310
320 | 48.1595 | 24.1711 | 320
330 | 47.5667 | 24.1685 | 330
340 | 47.8819 | 24.1558 | 340
350 | 48.1454 | 24.1523 | 350
360 | 47.9572 | 24.1452 | 360
370 | 48.1077 | 24.1416 | 370
380 | 48.2818 | 24.1346 | 380
390 | 48.1971 | 24.1296 | 390
400 | 47.8631 | 24.1276 | 400
410 | 48.2206 | 24.1199 | 410
420 | 47.9901 | 24.1154 | 420
430 | 48.6958 | 24.1133 | 430
440 | 47.642 | 24.1088 | 440
450 | 47.7361 | 24.1025 | 450
460 | 47.8302 | 24.0949 | 460
470 | 48.0889 | 24.0873 | 470
480 | 48.2442 | 24.088 | 480
490 | 47.9054 | 24.0819 | 490
500 | 47.9572 | 24.0778 | 500
510 | 48.0748 | 24.0758 | 510
520 | 48.0136 | 24.0673 | 520
530 | 48.0936 | 24.0695 | 530
540 | 48.1689 | 24.0606 | 540
550 | 48.1454 | 24.058 | 550
560 | 48.2583 | 24.0503 | 560
570 | 48.39 | 24.046 | 570
580 | 47.8114 | 24.0454 | 580
590 | 48.4088 | 24.0365 | 590
600 | 48.0136 | 24.0395 | 600
610 | 47.6185 | 24.0314 | 610
620 | 48.4041 | 24.0287 | 620
630 | 48.39 | 24.0236 | 630
640 | 48.3947 | 24.0194 | 640
650 | 48.1877 | 24.0191 | 650
660 | 48.2206 | 24.0126 | 660
670 | 48.7899 | 24.0119 | 670
680 | 47.9619 | 24.0057 | 680
690 | 48.0513 | 24.0044 | 690
700 | 48.3759 | 24.0038 | 700
710 | 48.5076 | 23.9969 | 710
720 | 48.3147 | 23.996 | 720
730 | 48.1454 | 23.9889 | 730
740 | 47.9619 | 23.9883 | 740
750 | 47.9196 | 23.9849 | 750
760 | 47.7549 | 23.981 | 760
770 | 48.1689 | 23.9785 | 770
780 | 48.0748 | 23.977 | 780
790 | 47.6185 | 23.972 | 790
800 | 48.0889 | 23.9699 | 800
810 | 48.0748 | 23.9676 | 810
820 | 48.1077 | 23.9666 | 820
830 | 47.9572 | 23.961 | 830
840 | 47.5997 | 23.9611 | 840
850 | 47.976 | 23.9604 | 850
860 | 47.6044 | 23.9515 | 860
870 | 47.6938 | 23.9515 | 870
880 | 48.2583 | 23.9494 | 880
890 | 47.7361 | 23.9529 | 890
900 | 48.0184 | 23.9417 | 900
910 | 48.4135 | 23.9413 | 910
920 | 48.5217 | 23.9361 | 920
930 | 47.9572 | 23.9359 | 930
940 | 47.9995 | 23.9352 | 940
950 | 48.2395 | 23.932 | 950
960 | 48.1454 | 23.9298 | 960
970 | 48.3524 | 23.9279 | 970
980 | 48.0372 | 23.926 | 980
990 | 48.0372 | 23.9252 | 990
1000 | 48.2583 | 23.9241 | 1000
1010 | 47.976 | 23.9213 | 1010
1020 | 48.5076 | 23.917 | 1020
1030 | 47.8067 | 23.9147 | 1030
1040 | 48.4888 | 23.9097 | 1040
1050 | 47.9995 | 23.9123 | 1050
1060 | 48.4653 | 23.9095 | 1060
1070 | 48.4088 | 23.912 | 1070
1080 | 48.5076 | 23.9087 | 1080
1090 | 48.2959 | 23.909 | 1090
1100 | 48.3147 | 23.9027 | 1100
1110 | 48.5641 | 23.9028 | 1110
1120 | 48.2442 | 23.8975 | 1120
1130 | 48.2395 | 23.9001 | 1130
1140 | 48.3759 | 23.8966 | 1140
1150 | 47.9243 | 23.8951 | 1150
1160 | 48.1454 | 23.8963 | 1160
1170 | 47.7502 | 23.8952 | 1170
1180 | 48.0936 | 23.8922 | 1180
1190 | 47.8114 | 23.8878 | 1190
1200 | 48.3524 | 23.8894 | 1200
1210 | 48.1689 | 23.8909 | 1210
1220 | 47.802 | 23.8836 | 1220
1230 | 48.1877 | 23.8891 | 1230
1240 | 47.9243 | 23.8825 | 1240
1250 | 48.0842 | 23.8807 | 1250
1260 | 47.6938 | 23.8785 | 1260
1270 | 47.9901 | 23.8767 | 1270
1280 | 48.5641 | 23.8777 | 1280
1290 | 47.8631 | 23.8736 | 1290
1300 | 48.5452 | 23.872 | 1300
1310 | 48.0372 | 23.8724 | 1310
1320 | 47.976 | 23.8755 | 1320
1330 | 47.8866 | 23.8736 | 1330
1340 | 47.9384 | 23.8712 | 1340
1350 | 48.0936 | 23.8672 | 1350
1360 | 48.1313 | 23.8709 | 1360
1370 | 48.1454 | 23.8725 | 1370
1380 | 48.5782 | 23.8674 | 1380
1390 | 48.3759 | 23.8675 | 1390
1400 | 48.1877 | 23.8621 | 1400
1410 | 47.5856 | 23.8687 | 1410
1420 | 47.9007 | 23.8626 | 1420
1430 | 48.3335 | 23.8628 | 1430
1440 | 47.9007 | 23.8588 | 1440
1450 | 47.802 | 23.8597 | 1450
1460 | 48.2771 | 23.8596 | 1460
1470 | 48.39 | 23.8578 | 1470
1480 | 47.8302 | 23.8525 | 1480
1490 | 48.3147 | 23.8545 | 1490
1500 | 48.6346 | 23.8516 | 1500
1510 | 48.2771 | 23.8519 | 1510
1520 | 48.1501 | 23.8543 | 1520
1530 | 47.8631 | 23.8497 | 1530
1540 | 47.802 | 23.8451 | 1540
1550 | 48.0513 | 23.8481 | 1550
1560 | 48.4512 | 23.8443 | 1560
1570 | 47.7737 | 23.8452 | 1570
1580 | 48.263 | 23.8418 | 1580
1590 | 48.3382 | 23.8431 | 1590
1600 | 48.0184 | 23.8427 | 1600
1610 | 48.0748 | 23.8461 | 1610
1620 | 47.8819 | 23.8412 | 1620
1630 | 48.0325 | 23.8427 | 1630
1640 | 48.263 | 23.8427 | 1640
1650 | 48.4135 | 23.838 | 1650
1660 | 48.7334 | 23.8405 | 1660
1670 | 47.9807 | 23.8374 | 1670
1680 | 47.9619 | 23.8405 | 1680
1690 | 48.3477 | 23.8382 | 1690
1700 | 47.9619 | 23.8344 | 1700
1710 | 48.263 | 23.8348 | 1710
1720 | 47.7361 | 23.8333 | 1720
1730 | 48.47 | 23.835 | 1730
1740 | 47.8443 | 23.8305 | 1740
1750 | 48.1077 | 23.8313 | 1750
1760 | 48.0325 | 23.8333 | 1760
1770 | 48.1689 | 23.8331 | 1770
1780 | 48.1642 | 23.8357 | 1780
1790 | 48.6911 | 23.8282 | 1790
1800 | 47.8866 | 23.8317 | 1800
1810 | 47.7126 | 23.8312 | 1810
1820 | 48.263 | 23.8276 | 1820
1830 | 48.2583 | 23.8292 | 1830
1840 | 47.8631 | 23.8231 | 1840
1850 | 48.4841 | 23.8271 | 1850
1860 | 48.0748 | 23.826 | 1860
1870 | 48.3524 | 23.8299 | 1870
1880 | 48.0184 | 23.8226 | 1880
1890 | 48.39 | 23.825 | 1890
1900 | 47.9995 | 23.8211 | 1900
1910 | 47.8584 | 23.8238 | 1910
1920 | 48.2818 | 23.8236 | 1920
1930 | 48.0089 | 23.8215 | 1930
1940 | 48.1077 | 23.8205 | 1940
1950 | 47.8678 | 23.8176 | 1950
1960 | 48.39 | 23.8214 | 1960
1970 | 47.9054 | 23.8174 | 1970
1980 | 47.9572 | 23.8193 | 1980
1990 | 48.2442 | 23.8182 | 1990
2000 | 47.9995 | 23.8252 | 2000
2010 | 47.9384 | 23.8201 | 2010
2020 | 48.3524 | 23.8164 | 2020
2030 | 48.3335 | 23.8198 | 2030
2040 | 47.9243 | 23.8136 | 2040
2050 | 47.9619 | 23.8182 | 2050
2060 | 47.8584 | 23.8172 | 2060
2070 | 48.4464 | 23.8203 | 2070
2080 | 48.5452 | 23.8117 | 2080
2090 | 48.1077 | 23.8166 | 2090
2100 | 48.0513 | 23.8155 | 2100
2110 | 47.9007 | 23.8116 | 2110
2120 | 48.39 | 23.811 | 2120
2130 | 48.263 | 23.8088 | 2130
2140 | 47.6796 | 23.8134 | 2140
2150 | 48.1124 | 23.8131 | 2150
2160 | 48.1501 | 23.808 | 2160
2170 | 48.5594 | 23.8092 | 2170
2180 | 47.9619 | 23.8099 | 2180
2190 | 48.1077 | 23.8057 | 2190
2200 | 47.8443 | 23.8125 | 2200
2210 | 47.9995 | 23.8071 | 2210
2220 | 48.2018 | 23.8099 | 2220
2230 | 48.263 | 23.8101 | 2230
2240 | 47.8114 | 23.8048 | 2240
2250 | 48.3382 | 23.8036 | 2250
2260 | 48.3194 | 23.8069 | 2260
2270 | 48.0936 | 23.8083 | 2270
2280 | 47.6796 | 23.8053 | 2280
2290 | 48.056 | 23.8099 | 2290
2300 | 47.7925 | 23.804 | 2300
2310 | 47.8302 | 23.8078 | 2310
2320 | 48.2253 | 23.8068 | 2320
2330 | 48.1689 | 23.8086 | 2330
2340 | 48.2253 | 23.8029 | 2340
2350 | 48.0889 | 23.8072 | 2350
2360 | 48.4088 | 23.8059 | 2360
2370 | 47.9007 | 23.8028 | 2370
2380 | 48.2253 | 23.8012 | 2380
2390 | 48.597 | 23.8026 | 2390
2400 | 48.1313 | 23.8022 | 2400
2410 | 47.8302 | 23.7996 | 2410
2420 | 48.4135 | 23.804 | 2420
2430 | 48.4888 | 23.8012 | 2430
2440 | 47.9807 | 23.8038 | 2440
2450 | 48.0325 | 23.8094 | 2450
2460 | 48.0889 | 23.8027 | 2460
2470 | 48.183 | 23.7997 | 2470
2480 | 47.8631 | 23.801 | 2480
2490 | 48.3382 | 23.7976 | 2490
2500 | 47.9431 | 23.8021 | 2500
2510 | 47.8866 | 23.7962 | 2510
2520 | 48.2818 | 23.7982 | 2520
2530 | 47.7831 | 23.8023 | 2530
2540 | 48.2395 | 23.7975 | 2540
2550 | 48.1124 | 23.8019 | 2550
2560 | 47.8819 | 23.8037 | 2560
2570 | 48.263 | 23.8045 | 2570
2580 | 48.0513 | 23.7925 | 2580
2590 | 48.0466 | 23.7974 | 2590
2600 | 48.1501 | 23.8043 | 2600
2610 | 48.39 | 23.7959 | 2610
2620 | 47.9007 | 23.7934 | 2620
2630 | 48.2253 | 23.7972 | 2630
2640 | 48.4135 | 23.8026 | 2640
2650 | 48.0513 | 23.7927 | 2650
2660 | 48.1454 | 23.7941 | 2660
2670 | 47.849 | 23.7955 | 2670
2680 | 47.5479 | 23.7972 | 2680
2690 | 48.597 | 23.7908 | 2690
2700 | 48.1266 | 23.7887 | 2700
2710 | 48.0136 | 23.7955 | 2710
2720 | 48.056 | 23.795 | 2720
2730 | 47.4679 | 23.7912 | 2730
2740 | 48.4512 | 23.7927 | 2740
2750 | 47.8866 | 23.7962 | 2750
2760 | 48.183 | 23.7962 | 2760
2770 | 48.0513 | 23.7894 | 2770
2780 | 47.8866 | 23.7959 | 2780
2790 | 47.9149 | 23.7924 | 2790
2800 | 47.6796 | 23.7955 | 2800
2810 | 47.7126 | 23.8003 | 2810
2820 | 48.3006 | 23.7878 | 2820
2830 | 48.3147 | 23.7877 | 2830
2840 | 47.9572 | 23.788 | 2840
2850 | 47.7173 | 23.793 | 2850
2860 | 47.9948 | 23.7868 | 2860
2870 | 47.9713 | 23.7868 | 2870
2880 | 47.8866 | 23.7857 | 2880
2890 | 48.2206 | 23.7917 | 2890
2900 | 48.0936 | 23.7911 | 2900
2910 | 48.4464 | 23.7875 | 2910
2920 | 48.1266 | 23.7895 | 2920
2930 | 48.1124 | 23.7876 | 2930
2940 | 48.3759 | 23.7902 | 2940
2950 | 48.2959 | 23.7863 | 2950
2960 | 48.0325 | 23.7884 | 2960
2970 | 48.2253 | 23.7865 | 2970
2980 | 47.435 | 23.7851 | 2980
2990 | 47.7502 | 23.7904 | 2990
3000 | 47.9243 | 23.784 | 3000
3010 | 48.0748 | 23.7899 | 3010
3020 | 48.4041 | 23.7928 | 3020
3030 | 48.0136 | 23.7875 | 3030
3040 | 48.2583 | 23.7867 | 3040
3050 | 47.8631 | 23.7845 | 3050
3060 | 48.0654 | 23.7844 | 3060
3070 | 48.2348 | 23.7939 | 3070
3080 | 48.3194 | 23.7877 | 3080
3090 | 47.9054 | 23.7917 | 3090
3100 | 48.0748 | 23.7867 | 3100
3110 | 48.6346 | 23.7832 | 3110
3120 | 48.3194 | 23.7831 | 3120
3130 | 47.8255 | 23.7925 | 3130
3140 | 48.1454 | 23.7889 | 3140
3150 | 47.7173 | 23.7836 | 3150
3160 | 48.0372 | 23.7844 | 3160
3170 | 48.056 | 23.7863 | 3170
3180 | 48.1642 | 23.7827 | 3180
3190 | 47.9525 | 23.7838 | 3190
3200 | 48.1689 | 23.784 | 3200
3210 | 48.2583 | 23.7796 | 3210
3220 | 48.1454 | 23.7801 | 3220
3230 | 47.9807 | 23.7804 | 3230
3240 | 47.8866 | 23.7809 | 3240
3250 | 48.3194 | 23.7808 | 3250
3260 | 48.0466 | 23.7837 | 3260
3270 | 48.3147 | 23.7806 | 3270
3280 | 48.4135 | 23.7844 | 3280
3290 | 48.2253 | 23.7822 | 3290
3300 | 48.0889 | 23.7816 | 3300
3310 | 48.0513 | 23.7792 | 3310
3320 | 48.6393 | 23.7858 | 3320
3330 | 48.0889 | 23.7856 | 3330
3340 | 47.8443 | 23.7816 | 3340
3350 | 47.8866 | 23.7785 | 3350
3360 | 48.3477 | 23.779 | 3360
3370 | 48.2442 | 23.7763 | 3370
3380 | 48.3335 | 23.7792 | 3380
3390 | 48.5641 | 23.7779 | 3390
3400 | 47.9572 | 23.7794 | 3400
3410 | 48.1689 | 23.7778 | 3410
3420 | 48.2065 | 23.7837 | 3420
3430 | 48.2395 | 23.7768 | 3430
3440 | 48.2442 | 23.7798 | 3440
3450 | 47.8819 | 23.7846 | 3450
3460 | 48.4464 | 23.7804 | 3460
3470 | 48.5452 | 23.7788 | 3470
3480 | 48.4653 | 23.7814 | 3480
3490 | 47.9007 | 23.7828 | 3490
3500 | 48.0748 | 23.7774 | 3500
3510 | 48.39 | 23.7863 | 3510
3520 | 48.2253 | 23.7803 | 3520
3530 | 47.7314 | 23.7794 | 3530
3540 | 48.4888 | 23.7746 | 3540
3550 | 47.5667 | 23.7789 | 3550
3560 | 47.9619 | 23.7776 | 3560
3570 | 48.5782 | 23.7776 | 3570
3580 | 48.4135 | 23.7762 | 3580
3590 | 48.4464 | 23.7711 | 3590
3600 | 48.0513 | 23.774 | 3600
3610 | 47.9525 | 23.7765 | 3610
3620 | 47.9243 | 23.7776 | 3620
3630 | 48.4888 | 23.7761 | 3630
3640 | 48.1689 | 23.7745 | 3640
3650 | 47.6608 | 23.7759 | 3650
3660 | 48.3335 | 23.774 | 3660
3670 | 48.5217 | 23.7734 | 3670
3680 | 48.0936 | 23.7764 | 3680
3690 | 48.0136 | 23.7735 | 3690
3700 | 48.0513 | 23.7753 | 3700
3710 | 48.3759 | 23.7771 | 3710
3720 | 47.7878 | 23.7768 | 3720
3730 | 48.0513 | 23.7784 | 3730
3740 | 47.9572 | 23.7761 | 3740
3750 | 47.9525 | 23.7738 | 3750
3760 | 48.3524 | 23.7751 | 3760
3770 | 48.2253 | 23.7747 | 3770
3780 | 48.3382 | 23.7713 | 3780
3790 | 47.7361 | 23.7743 | 3790
3800 | 48.47 | 23.7822 | 3800
3810 | 48.3194 | 23.7726 | 3810
3820 | 47.7126 | 23.7743 | 3820
3830 | 48.2018 | 23.774 | 3830
3840 | 48.6581 | 23.7743 | 3840
3850 | 47.8866 | 23.7779 | 3850
3860 | 48.1642 | 23.771 | 3860
3870 | 48.4135 | 23.7747 | 3870
3880 | 48.0089 | 23.7685 | 3880
3890 | 48.4135 | 23.7727 | 3890
3900 | 48.1877 | 23.7716 | 3900
3910 | 48.4276 | 23.7715 | 3910
3920 | 47.7361 | 23.778 | 3920
3930 | 48.0701 | 23.7736 | 3930
3940 | 47.9572 | 23.7761 | 3940
3950 | 48.3524 | 23.7711 | 3950
3960 | 48.056 | 23.7738 | 3960
3970 | 48.3477 | 23.7717 | 3970
3980 | 47.9243 | 23.7749 | 3980
3990 | 48.3712 | 23.7763 | 3990
4000 | 48.0748 | 23.7719 | 4000
4010 | 48.0889 | 23.7719 | 4010
4020 | 47.8302 | 23.7723 | 4020
4030 | 47.8067 | 23.7755 | 4030
4040 | 47.6749 | 23.7687 | 4040
4050 | 47.8631 | 23.7738 | 4050
4060 | 47.769 | 23.7714 | 4060
4070 | 48.597 | 23.7716 | 4070
4080 | 48.3194 | 23.7708 | 4080
4090 | 47.9572 | 23.7736 | 4090
4100 | 48.0889 | 23.7732 | 4100
4110 | 48.0513 | 23.7722 | 4110
4120 | 48.1077 | 23.7701 | 4120
4130 | 47.9384 | 23.7742 | 4130
4140 | 48.0136 | 23.77 | 4140
4150 | 48.3006 | 23.7689 | 4150
4160 | 47.9243 | 23.7719 | 4160
4170 | 47.5997 | 23.7647 | 4170
4180 | 48.0748 | 23.7672 | 4180
4190 | 47.9431 | 23.7677 | 4190
4200 | 48.1642 | 23.7749 | 4200
4210 | 48.263 | 23.7682 | 4210
4220 | 48.1689 | 23.7677 | 4220
4230 | 48.4135 | 23.7718 | 4230
4240 | 48.056 | 23.7713 | 4240
4250 | 48.263 | 23.7669 | 4250
4260 | 47.9243 | 23.7687 | 4260
4270 | 47.6608 | 23.7701 | 4270
4280 | 47.849 | 23.7716 | 4280
4290 | 48.3712 | 23.7702 | 4290
4300 | 48.1689 | 23.7714 | 4300
4310 | 48.2253 | 23.7682 | 4310
4320 | 47.9807 | 23.7687 | 4320
4330 | 47.7878 | 23.7741 | 4330
4340 | 47.9619 | 23.7767 | 4340
4350 | 47.9337 | 23.7676 | 4350
4360 | 48.3759 | 23.7671 | 4360
4370 | 48.3147 | 23.7715 | 4370
4380 | 47.9149 | 23.7666 | 4380
4390 | 47.9431 | 23.7706 | 4390
4400 | 48.1077 | 23.7703 | 4400
4410 | 48.1689 | 23.7657 | 4410
4420 | 48.5076 | 23.7629 | 4420
4430 | 48.39 | 23.7619 | 4430
4440 | 47.8819 | 23.7675 | 4440
4450 | 48.0889 | 23.7641 | 4450
4460 | 47.9572 | 23.7694 | 4460
4470 | 48.1124 | 23.7735 | 4470
4480 | 48.0513 | 23.768 | 4480
4490 | 48.4888 | 23.7636 | 4490
4500 | 47.9713 | 23.7689 | 4500
4510 | 48.0184 | 23.7686 | 4510
4520 | 48.771 | 23.7698 | 4520
4530 | 47.976 | 23.7658 | 4530
4540 | 48.263 | 23.7696 | 4540
4550 | 47.769 | 23.766 | 4550
4560 | 48.1454 | 23.7633 | 4560
4570 | 48.2818 | 23.7672 | 4570
4580 | 47.5856 | 23.7643 | 4580
4590 | 48.2959 | 23.7737 | 4590
4600 | 48.0701 | 23.7655 | 4600
4610 | 47.9619 | 23.7698 | 4610
4620 | 47.9807 | 23.768 | 4620
4630 | 47.8302 | 23.7657 | 4630
4640 | 47.8866 | 23.7687 | 4640
4650 | 48.2253 | 23.7678 | 4650
4660 | 47.8067 | 23.7639 | 4660
4670 | 48.1407 | 23.7702 | 4670
4680 | 48.1501 | 23.7656 | 4680
4690 | 48.2959 | 23.766 | 4690
4700 | 48.39 | 23.7611 | 4700
4710 | 47.8631 | 23.7681 | 4710
4720 | 48.3759 | 23.7654 | 4720
4730 | 47.642 | 23.7613 | 4730
4740 | 48.1454 | 23.7616 | 4740
4750 | 48.0278 | 23.7602 | 4750
4760 | 48.0136 | 23.7674 | 4760
4770 | 47.9572 | 23.7647 | 4770
4780 | 47.6326 | 23.7634 | 4780
4790 | 47.9572 | 23.7682 | 4790
4800 | 47.5808 | 23.7658 | 4800
4810 | 48.2442 | 23.7644 | 4810
4820 | 48.3571 | 23.7599 | 4820
4830 | 48.3382 | 23.7602 | 4830
4840 | 48.7522 | 23.7641 | 4840
4850 | 47.9431 | 23.7704 | 4850
4860 | 48.0372 | 23.763 | 4860
4870 | 48.263 | 23.7627 | 4870
4880 | 47.562 | 23.7654 | 4880
4890 | 48.1877 | 23.7592 | 4890
4900 | 48.1313 | 23.7652 | 4900
4910 | 47.9807 | 23.762 | 4910
4920 | 48.0513 | 23.7635 | 4920
4930 | 48.0513 | 23.7621 | 4930
4940 | 47.8302 | 23.768 | 4940
4950 | 47.9431 | 23.7628 | 4950
4960 | 48.6346 | 23.7644 | 4960
4970 | 48.2818 | 23.7606 | 4970
4980 | 48.1877 | 23.7653 | 4980
4990 | 47.8396 | 23.7634 | 4990
5000 | 47.9384 | 23.7645 | 5000
5010 | 48.2065 | 23.7657 | 5010
5020 | 48.4841 | 23.7623 | 5020
5030 | 47.9995 | 23.7625 | 5030
5040 | 48.3194 | 23.761 | 5040
5050 | 48.2018 | 23.7635 | 5050
5060 | 47.8631 | 23.7613 | 5060
5070 | 47.9619 | 23.7598 | 5070
5080 | 48.0889 | 23.7611 | 5080
5090 | 48.3524 | 23.7694 | 5090
5100 | 48.0936 | 23.7591 | 5100
5110 | 47.9007 | 23.7597 | 5110
5120 | 48.1877 | 23.767 | 5120
5130 | 48.2818 | 23.7635 | 5130
5140 | 48.2018 | 23.7635 | 5140
5150 | 48.0889 | 23.7659 | 5150
5160 | 47.9054 | 23.7632 | 5160
5170 | 48.2771 | 23.7629 | 5170
5180 | 48.4135 | 23.7604 | 5180
5190 | 48.1266 | 23.7614 | 5190
5200 | 48.2818 | 23.7633 | 5200
5210 | 48.0325 | 23.7614 | 5210
5220 | 48.0184 | 23.7579 | 5220
5230 | 47.7878 | 23.7608 | 5230
5240 | 47.8443 | 23.7599 | 5240
5250 | 48.2724 | 23.7637 | 5250
5260 | 48.1124 | 23.7571 | 5260
5270 | 47.9807 | 23.7613 | 5270
5280 | 48.3147 | 23.7635 | 5280
5290 | 47.9243 | 23.7622 | 5290
5300 | 48.0136 | 23.7579 | 5300
5310 | 47.976 | 23.7577 | 5310
5320 | 47.7549 | 23.7595 | 5320
5330 | 47.9243 | 23.7589 | 5330
5340 | 48.597 | 23.7577 | 5340
5350 | 48.2959 | 23.7569 | 5350
5360 | 48.0842 | 23.7627 | 5360
5370 | 48.0748 | 23.76 | 5370
5380 | 48.0748 | 23.7587 | 5380
5390 | 48.3759 | 23.7556 | 5390
5400 | 48.0748 | 23.7597 | 5400
5410 | 47.6185 | 23.7674 | 5410
5420 | 48.0136 | 23.7578 | 5420
5430 | 48.0701 | 23.7624 | 5430
5440 | 47.4679 | 23.7652 | 5440
5450 | 48.3571 | 23.7598 | 5450
5460 | 48.0513 | 23.7559 | 5460
5470 | 48.2018 | 23.7622 | 5470
5480 | 48.0936 | 23.7629 | 5480
5490 | 48.3194 | 23.7589 | 5490
5500 | 47.7126 | 23.7577 | 5500
5510 | 48.3571 | 23.7644 | 5510
5520 | 48.0184 | 23.754 | 5520
5530 | 48.1689 | 23.7608 | 5530
5540 | 48.31 | 23.7547 | 5540
5550 | 48.0748 | 23.7546 | 5550
5560 | 47.8866 | 23.7568 | 5560
5570 | 48.0513 | 23.7562 | 5570
5580 | 48.3382 | 23.7539 | 5580
5590 | 47.8866 | 23.7551 | 5590
5600 | 48.0748 | 23.7573 | 5600
5610 | 47.8302 | 23.7568 | 5610
5620 | 47.7126 | 23.7617 | 5620
5630 | 47.9243 | 23.7581 | 5630
5640 | 47.8443 | 23.7552 | 5640
5650 | 48.2583 | 23.7603 | 5650
5660 | 47.8302 | 23.7589 | 5660
5670 | 48.2018 | 23.7586 | 5670
5680 | 47.9807 | 23.7596 | 5680
5690 | 48.1266 | 23.7577 | 5690
5700 | 48.2583 | 23.7619 | 5700
5710 | 48.0466 | 23.7616 | 5710
5720 | 48.2206 | 23.7638 | 5720
5730 | 48.103 | 23.7611 | 5730
5740 | 48.1313 | 23.7582 | 5740
5750 | 48.5405 | 23.762 | 5750
5760 | 48.1077 | 23.758 | 5760
5770 | 47.9619 | 23.7628 | 5770
5780 | 48.103 | 23.7593 | 5780
5790 | 47.9431 | 23.7616 | 5790
5800 | 47.7549 | 23.7568 | 5800
5810 | 47.6514 | 23.7585 | 5810
5820 | 47.769 | 23.761 | 5820
5830 | 47.6796 | 23.7608 | 5830
5840 | 47.9431 | 23.7593 | 5840
5850 | 48.5217 | 23.7609 | 5850
5860 | 48.2018 | 23.7594 | 5860
5870 | 48.2018 | 23.7564 | 5870
5880 | 48.1266 | 23.7561 | 5880
5890 | 48.0513 | 23.7599 | 5890
5900 | 48.263 | 23.7605 | 5900
5910 | 47.9619 | 23.7534 | 5910
5920 | 48.0513 | 23.7575 | 5920
5930 | 48.0136 | 23.7577 | 5930
5940 | 48.39 | 23.7595 | 5940
5950 | 47.7314 | 23.7571 | 5950
5960 | 47.9525 | 23.7604 | 5960
5970 | 48.1454 | 23.7617 | 5970
5980 | 48.1077 | 23.7582 | 5980
5990 | 47.9243 | 23.7553 | 5990
6000 | 48.1313 | 23.7542 | 6000
6010 | 47.6185 | 23.7548 | 6010
6020 | 48.1407 | 23.7518 | 6020
6030 | 48.3712 | 23.7533 | 6030
6040 | 48.39 | 23.753 | 6040
6050 | 48.0936 | 23.7588 | 6050
6060 | 47.7126 | 23.7585 | 6060
6070 | 47.9995 | 23.7559 | 6070
6080 | 47.8866 | 23.7564 | 6080
6090 | 48.7663 | 23.7559 | 6090
6100 | 47.5479 | 23.7594 | 6100
6110 | 47.9337 | 23.7567 | 6110
6120 | 48.0136 | 23.7597 | 6120
6130 | 47.7361 | 23.7548 | 6130
6140 | 47.8255 | 23.7564 | 6140
6150 | 47.8678 | 23.7537 | 6150
6160 | 48.3947 | 23.7574 | 6160
6170 | 47.8631 | 23.7596 | 6170
6180 | 48.0889 | 23.7537 | 6180
6190 | 48.0184 | 23.7526 | 6190
6200 | 47.7878 | 23.7564 | 6200
6210 | 48.7146 | 23.7564 | 6210
6220 | 47.8114 | 23.7541 | 6220
6230 | 47.8678 | 23.7547 | 6230
6240 | 47.9243 | 23.7565 | 6240
6250 | 47.8631 | 23.7559 | 6250
6260 | 47.7925 | 23.7529 | 6260
6270 | 48.2018 | 23.7591 | 6270
6280 | 47.9619 | 23.7526 | 6280
6290 | 48.0372 | 23.7589 | 6290
6300 | 47.9243 | 23.7561 | 6300
6310 | 48.2771 | 23.7538 | 6310
6320 | 47.9007 | 23.7621 | 6320
6330 | 47.9054 | 23.7575 | 6330
6340 | 47.6985 | 23.7588 | 6340
6350 | 47.642 | 23.7559 | 6350
6360 | 48.0513 | 23.7556 | 6360
6370 | 48.0325 | 23.7589 | 6370
6380 | 47.7878 | 23.7549 | 6380
6390 | 48.1077 | 23.7555 | 6390
6400 | 47.8255 | 23.761 | 6400
6410 | 48.0748 | 23.7587 | 6410
6420 | 48.2159 | 23.7525 | 6420
6430 | 47.8067 | 23.7557 | 6430
6440 | 47.5856 | 23.7548 | 6440
6450 | 47.849 | 23.7564 | 6450
6460 | 48.0372 | 23.7515 | 6460
6470 | 47.7361 | 23.7528 | 6470
6480 | 48.2771 | 23.7525 | 6480
6490 | 47.6796 | 23.7541 | 6490
6500 | 48.1454 | 23.7543 | 6500
6510 | 48.2583 | 23.7548 | 6510
6520 | 47.976 | 23.7558 | 6520
6530 | 48.2018 | 23.7577 | 6530
6540 | 47.6938 | 23.7535 | 6540
6550 | 48.0889 | 23.7616 | 6550
6560 | 48.3524 | 23.7519 | 6560
6570 | 48.3194 | 23.7535 | 6570
6580 | 48.1077 | 23.7594 | 6580
6590 | 47.6985 | 23.7532 | 6590
6600 | 47.9572 | 23.7563 | 6600
6610 | 47.9619 | 23.7541 | 6610
6620 | 48.5641 | 23.7518 | 6620
6630 | 48.597 | 23.7557 | 6630
6640 | 48.056 | 23.7554 | 6640
6650 | 48.263 | 23.7573 | 6650
6660 | 47.9431 | 23.7544 | 6660
6670 | 47.6985 | 23.7528 | 6670
6680 | 48.183 | 23.7583 | 6680
6690 | 48.183 | 23.753 | 6690
6700 | 48.0889 | 23.7514 | 6700
6710 | 47.8255 | 23.7529 | 6710
6720 | 48.3006 | 23.7569 | 6720
6730 | 47.8302 | 23.7536 | 6730
6740 | 47.9619 | 23.7533 | 6740
6750 | 47.9995 | 23.7528 | 6750
6760 | 47.9431 | 23.754 | 6760
6770 | 47.6185 | 23.7539 | 6770
6780 | 47.849 | 23.7514 | 6780
6790 | 48.3147 | 23.7535 | 6790
6800 | 48.0701 | 23.7524 | 6800
6810 | 48.1124 | 23.758 | 6810
6820 | 48.1077 | 23.7507 | 6820
6830 | 48.2159 | 23.7552 | 6830
6840 | 47.9995 | 23.7523 | 6840
6850 | 48.2442 | 23.7484 | 6850
6860 | 47.9196 | 23.7489 | 6860
6870 | 48.0184 | 23.7563 | 6870
6880 | 47.7549 | 23.7493 | 6880
6890 | 48.0513 | 23.751 | 6890
6900 | 48.0513 | 23.7557 | 6900
6910 | 47.6232 | 23.753 | 6910
6920 | 48.0184 | 23.7524 | 6920
6930 | 47.5103 | 23.754 | 6930
6940 | 47.8866 | 23.751 | 6940
6950 | 48.2959 | 23.7554 | 6950
6960 | 48.2018 | 23.7561 | 6960
6970 | 48.1642 | 23.7515 | 6970
6980 | 47.7126 | 23.7552 | 6980
6990 | 48.0889 | 23.75 | 6990
7000 | 47.8866 | 23.7537 | 7000
7010 | 48.263 | 23.7567 | 7010
7020 | 48.2065 | 23.7539 | 7020
7030 | 47.7126 | 23.7485 | 7030
7040 | 48.3524 | 23.7537 | 7040
7050 | 48.2583 | 23.7519 | 7050
7060 | 47.7361 | 23.7583 | 7060
7070 | 47.6044 | 23.7543 | 7070
7080 | 48.2959 | 23.7509 | 7080
7090 | 47.6938 | 23.7524 | 7090
7100 | 48.0325 | 23.7528 | 7100
7110 | 48.103 | 23.754 | 7110
7120 | 48.0136 | 23.7507 | 7120
7130 | 48.0889 | 23.7504 | 7130
7140 | 47.9243 | 23.7544 | 7140
7150 | 47.9384 | 23.7558 | 7150
7160 | 47.642 | 23.7461 | 7160
7170 | 48.0748 | 23.752 | 7170
7180 | 48.5829 | 23.7533 | 7180
7190 | 47.9007 | 23.7531 | 7190
7200 | 48.1077 | 23.751 | 7200
7210 | 47.7361 | 23.7507 | 7210
7220 | 47.8866 | 23.753 | 7220
7230 | 48.183 | 23.7488 | 7230
7240 | 48.0748 | 23.7522 | 7240
7250 | 48.0136 | 23.751 | 7250
7260 | 48.2442 | 23.7519 | 7260
7270 | 48.2253 | 23.7476 | 7270
7280 | 47.9619 | 23.7504 | 7280
7290 | 47.9007 | 23.7596 | 7290
7300 | 47.9948 | 23.7558 | 7300
7310 | 48.1124 | 23.7464 | 7310
7320 | 48.2959 | 23.7458 | 7320
7330 | 47.976 | 23.7493 | 7330
7340 | 48.2959 | 23.751 | 7340
7350 | 48.0748 | 23.7534 | 7350
7360 | 48.0372 | 23.7483 | 7360
7370 | 47.6232 | 23.7509 | 7370
7380 | 47.7126 | 23.7483 | 7380
7390 | 48.1501 | 23.7523 | 7390
7400 | 47.8067 | 23.7488 | 7400
7410 | 48.3524 | 23.7473 | 7410
7420 | 47.8255 | 23.7498 | 7420
7430 | 48.1077 | 23.7534 | 7430
7440 | 47.8631 | 23.7517 | 7440
7450 | 47.8302 | 23.7498 | 7450
7460 | 48.2018 | 23.7514 | 7460
7470 | 47.8678 | 23.7516 | 7470
7480 | 47.9619 | 23.7512 | 7480
7490 | 47.9243 | 23.7495 | 7490
7500 | 48.1266 | 23.7511 | 7500
7510 | 47.9196 | 23.7495 | 7510
7520 | 48.2253 | 23.7509 | 7520
7530 | 47.9807 | 23.7507 | 7530
7540 | 48.056 | 23.7481 | 7540
7550 | 48.1266 | 23.7494 | 7550
7560 | 47.562 | 23.7448 | 7560
7570 | 47.7878 | 23.7531 | 7570
7580 | 47.9054 | 23.7458 | 7580
7590 | 47.8631 | 23.745 | 7590
7600 | 47.8067 | 23.7482 | 7600
7610 | 47.9054 | 23.7475 | 7610
7620 | 47.9807 | 23.7508 | 7620
7630 | 47.8255 | 23.7511 | 7630
7640 | 48.2065 | 23.7506 | 7640
7650 | 48.0372 | 23.7457 | 7650
7660 | 47.9007 | 23.7479 | 7660
7670 | 47.9995 | 23.7516 | 7670
7680 | 48.1642 | 23.7487 | 7680
7690 | 48.3524 | 23.7526 | 7690
7700 | 48.1642 | 23.7475 | 7700
7710 | 48.0513 | 23.7477 | 7710
7720 | 47.849 | 23.75 | 7720
7730 | 48.2018 | 23.7482 | 7730
7740 | 48.1689 | 23.7517 | 7740
7750 | 47.7549 | 23.7454 | 7750
7760 | 48.2253 | 23.7487 | 7760
7770 | 48.0748 | 23.7488 | 7770
7780 | 47.8678 | 23.7467 | 7780
7790 | 48.2395 | 23.7495 | 7790
7800 | 48.263 | 23.7474 | 7800
7810 | 47.8866 | 23.7485 | 7810
7820 | 48.1689 | 23.7499 | 7820
7830 | 48.2206 | 23.751 | 7830
7840 | 48.2442 | 23.7477 | 7840
7850 | 48.0278 | 23.7492 | 7850
7860 | 47.9196 | 23.7515 | 7860
7870 | 47.9196 | 23.7458 | 7870
7880 | 47.8255 | 23.7474 | 7880
7890 | 47.9243 | 23.748 | 7890
7900 | 48.2065 | 23.7544 | 7900
7910 | 48.1689 | 23.7497 | 7910
7920 | 48.1266 | 23.7467 | 7920
7930 | 47.5667 | 23.7466 | 7930
7940 | 47.9995 | 23.7545 | 7940
7950 | 48.1877 | 23.7515 | 7950
7960 | 47.7173 | 23.7468 | 7960
7970 | 48.5217 | 23.7526 | 7970
7980 | 48.1077 | 23.7485 | 7980
7990 | 47.9995 | 23.7507 | 7990
8000 | 47.9054 | 23.7461 | 8000
8010 | 47.9948 | 23.7528 | 8010
8020 | 47.7361 | 23.7479 | 8020
8030 | 48.0748 | 23.7456 | 8030
8040 | 47.6796 | 23.7468 | 8040
8050 | 48.056 | 23.7477 | 8050
8060 | 48.47 | 23.7466 | 8060
8070 | 48.2018 | 23.7486 | 8070
8080 | 48.2253 | 23.746 | 8080
8090 | 47.5667 | 23.746 | 8090
8100 | 47.9572 | 23.7488 | 8100
8110 | 48.0136 | 23.7463 | 8110
8120 | 48.4464 | 23.7474 | 8120
8130 | 48.183 | 23.7489 | 8130
8140 | 48.1266 | 23.7458 | 8140
8150 | 48.3335 | 23.7516 | 8150
8160 | 47.9713 | 23.7495 | 8160
8170 | 48.2583 | 23.7528 | 8170
8180 | 48.5641 | 23.7516 | 8180
8190 | 48.0513 | 23.7487 | 8190
8200 | 47.8631 | 23.7503 | 8200
8210 | 47.9807 | 23.7511 | 8210
8220 | 47.6561 | 23.7453 | 8220
8230 | 48.1642 | 23.7519 | 8230
8240 | 48.0325 | 23.7447 | 8240
8250 | 48.2771 | 23.7507 | 8250
8260 | 48.2065 | 23.7498 | 8260
8270 | 48.0325 | 23.7472 | 8270
8280 | 48.4323 | 23.7442 | 8280
8290 | 48.1877 | 23.7455 | 8290
8300 | 48.2912 | 23.7438 | 8300
8310 | 48.4888 | 23.7432 | 8310
8320 | 48.1313 | 23.7477 | 8320
8330 | 47.8067 | 23.7532 | 8330
8340 | 48.1689 | 23.7434 | 8340
8350 | 48.0748 | 23.7428 | 8350
8360 | 48.1501 | 23.7472 | 8360
8370 | 48.3712 | 23.7519 | 8370
8380 | 47.9807 | 23.7421 | 8380
8390 | 48.263 | 23.7459 | 8390
8400 | 47.9572 | 23.7464 | 8400
8410 | 47.7549 | 23.7444 | 8410
8420 | 48.4276 | 23.7511 | 8420
8430 | 47.9572 | 23.7501 | 8430
8440 | 47.8866 | 23.7453 | 8440
8450 | 48.0513 | 23.7504 | 8450
8460 | 48.3194 | 23.7483 | 8460
8470 | 47.8114 | 23.7456 | 8470
8480 | 48.0889 | 23.7444 | 8480
8490 | 47.7361 | 23.7444 | 8490
8500 | 47.9807 | 23.7448 | 8500
8510 | 48.3382 | 23.7472 | 8510
8520 | 47.8302 | 23.7461 | 8520
8530 | 48.5594 | 23.7457 | 8530
8540 | 48.103 | 23.7522 | 8540
8550 | 47.769 | 23.7465 | 8550
8560 | 47.9525 | 23.7487 | 8560
8570 | 47.8631 | 23.7455 | 8570
8580 | 47.9948 | 23.7435 | 8580
8590 | 48.2959 | 23.749 | 8590
8600 | 48.1877 | 23.742 | 8600
8610 | 47.8067 | 23.7445 | 8610
8620 | 47.9572 | 23.7459 | 8620
8630 | 47.9525 | 23.7461 | 8630
8640 | 48.2065 | 23.7462 | 8640
8650 | 48.2583 | 23.7467 | 8650
8660 | 47.9243 | 23.7449 | 8660
8670 | 48.1689 | 23.741 | 8670
8680 | 47.9243 | 23.7477 | 8680
8690 | 48.0089 | 23.7451 | 8690
8700 | 48.4135 | 23.7489 | 8700
8710 | 47.8866 | 23.7486 | 8710
8720 | 47.7314 | 23.7441 | 8720
8730 | 48.0748 | 23.7472 | 8730
8740 | 48.263 | 23.7525 | 8740
8750 | 47.9572 | 23.7436 | 8750
8760 | 48.3382 | 23.7488 | 8760
8770 | 47.7173 | 23.7477 | 8770
8780 | 48.1877 | 23.7501 | 8780
8790 | 48.2818 | 23.7476 | 8790
8800 | 47.7173 | 23.7481 | 8800
8810 | 47.8866 | 23.7455 | 8810
8820 | 47.9243 | 23.7443 | 8820
8830 | 47.8631 | 23.748 | 8830
8840 | 48.2771 | 23.7486 | 8840
8850 | 48.0936 | 23.748 | 8850
8860 | 47.5244 | 23.7458 | 8860
8870 | 47.769 | 23.75 | 8870
8880 | 48.4135 | 23.7472 | 8880
8890 | 47.8678 | 23.747 | 8890
8900 | 48.2442 | 23.746 | 8900
8910 | 48.1077 | 23.7443 | 8910
8920 | 48.056 | 23.7498 | 8920
8930 | 47.9243 | 23.7492 | 8930
8940 | 47.769 | 23.7427 | 8940
8950 | 48.1642 | 23.7439 | 8950
8960 | 48.1124 | 23.7486 | 8960
8970 | 48.2206 | 23.7474 | 8970
8980 | 48.2959 | 23.7442 | 8980
8990 | 48.0889 | 23.744 | 8990
9000 | 48.0748 | 23.7494 | 9000
9010 | 48.3006 | 23.7438 | 9010
9020 | 48.2583 | 23.7441 | 9020
9030 | 47.8866 | 23.7447 | 9030
9040 | 48.2771 | 23.744 | 9040
9050 | 48.0889 | 23.7444 | 9050
9060 | 48.39 | 23.7461 | 9060
9070 | 47.8631 | 23.7479 | 9070
9080 | 47.8678 | 23.7414 | 9080
9090 | 48.3524 | 23.745 | 9090
9100 | 48.0184 | 23.7418 | 9100
9110 | 48.0889 | 23.7441 | 9110
9120 | 48.3335 | 23.7493 | 9120
9130 | 47.8866 | 23.743 | 9130
9140 | 48.4135 | 23.747 | 9140
9150 | 47.8631 | 23.7415 | 9150
9160 | 48.0748 | 23.747 | 9160
9170 | 47.8443 | 23.7426 | 9170
9180 | 48.263 | 23.7453 | 9180
9190 | 48.1454 | 23.75 | 9190
9200 | 47.8866 | 23.7439 | 9200
9210 | 48.3006 | 23.7453 | 9210
9220 | 47.9807 | 23.7439 | 9220
9230 | 48.3147 | 23.7455 | 9230
9240 | 48.0136 | 23.7507 | 9240
9250 | 47.8302 | 23.7457 | 9250
9260 | 47.8443 | 23.747 | 9260
9270 | 48.2771 | 23.7463 | 9270
9280 | 47.7878 | 23.7474 | 9280
9290 | 48.3147 | 23.7465 | 9290
9300 | 48.2395 | 23.7433 | 9300
9310 | 47.9525 | 23.7462 | 9310
9320 | 48.3147 | 23.7452 | 9320
9330 | 47.8772 | 23.7429 | 9330
9340 | 47.9948 | 23.7437 | 9340
9350 | 47.9243 | 23.7432 | 9350
9360 | 47.9007 | 23.7457 | 9360
9370 | 47.7361 | 23.7458 | 9370
9380 | 48.3524 | 23.7435 | 9380
9390 | 48.1642 | 23.746 | 9390
9400 | 48.3194 | 23.7395 | 9400
9410 | 47.689 | 23.7411 | 9410
9420 | 48.0748 | 23.7451 | 9420
9430 | 47.9572 | 23.7433 | 9430
9440 | 48.2018 | 23.7425 | 9440
9450 | 47.4915 | 23.7401 | 9450
9460 | 47.976 | 23.7387 | 9460
9470 | 48.39 | 23.7405 | 9470
9480 | 48.39 | 23.7397 | 9480
9490 | 48.3712 | 23.7478 | 9490
9500 | 47.8302 | 23.7413 | 9500
9510 | 48.0372 | 23.7456 | 9510
9520 | 47.9572 | 23.7415 | 9520
9530 | 48.0136 | 23.7414 | 9530
9540 | 48.5641 | 23.7445 | 9540
9550 | 47.9243 | 23.7422 | 9550
9560 | 47.9619 | 23.7422 | 9560
9570 | 47.8114 | 23.7415 | 9570
9580 | 47.9243 | 23.7448 | 9580
9590 | 48.3006 | 23.7458 | 9590
9600 | 47.6185 | 23.7418 | 9600
9610 | 47.8819 | 23.7404 | 9610
9620 | 48.2583 | 23.7396 | 9620
9630 | 48.1077 | 23.7462 | 9630
9640 | 48.47 | 23.7413 | 9640
9650 | 47.562 | 23.7505 | 9650
9660 | 47.849 | 23.7449 | 9660
9670 | 48.1454 | 23.7409 | 9670
9680 | 48.0748 | 23.7452 | 9680
9690 | 48.3147 | 23.7461 | 9690
9700 | 48.4088 | 23.7462 | 9700
9710 | 47.7173 | 23.7428 | 9710
9720 | 48.3194 | 23.7442 | 9720
9730 | 48.0513 | 23.7411 | 9730
9740 | 47.7502 | 23.7445 | 9740
9750 | 48.4088 | 23.7465 | 9750
9760 | 47.9948 | 23.7417 | 9760
9770 | 48.263 | 23.7528 | 9770
9780 | 48.2442 | 23.7474 | 9780
9790 | 47.9525 | 23.7433 | 9790
9800 | 47.8443 | 23.7415 | 9800
9810 | 48.0889 | 23.7411 | 9810
9820 | 48.0466 | 23.7474 | 9820
9830 | 47.769 | 23.7444 | 9830
9840 | 47.8067 | 23.746 | 9840
9850 | 47.8302 | 23.7447 | 9850
9860 | 48.1642 | 23.7404 | 9860
9870 | 47.9054 | 23.7463 | 9870
9880 | 47.9807 | 23.7397 | 9880
9890 | 48.3194 | 23.7429 | 9890
9900 | 47.9948 | 23.7426 | 9900
9910 | 48.1689 | 23.7422 | 9910
9920 | 48.5217 | 23.7433 | 9920
9930 | 47.9525 | 23.7412 | 9930
9940 | 48.1501 | 23.7461 | 9940
9950 | 48.183 | 23.7403 | 9950
9960 | 48.0184 | 23.7419 | 9960
9970 | 47.9431 | 23.7396 | 9970
9980 | 48.4088 | 23.7433 | 9980
9990 | 47.8866 | 23.7384 | 9990
10000 | 47.642 | 23.7428 | 10000
10010 | 48.5217 | 23.7468 | 10010
10020 | 48.1124 | 23.7434 | 10020
10030 | 48.1454 | 23.7427 | 10030
10040 | 47.8819 | 23.7416 | 10040
10050 | 48.39 | 23.7413 | 10050
10060 | 48.1689 | 23.7451 | 10060
10070 | 47.769 | 23.7403 | 10070
10080 | 48.4135 | 23.7437 | 10080
10090 | 48.2959 | 23.7467 | 10090
10100 | 48.0701 | 23.7412 | 10100
10110 | 47.8302 | 23.7433 | 10110
10120 | 48.3194 | 23.7381 | 10120
10130 | 47.9243 | 23.7452 | 10130
10140 | 48.0089 | 23.7418 | 10140
10150 | 47.9196 | 23.7406 | 10150
10160 | 47.9807 | 23.7473 | 10160
10170 | 47.8302 | 23.7455 | 10170
10180 | 47.9807 | 23.7444 | 10180
10190 | 48.3335 | 23.7421 | 10190
10200 | 47.9901 | 23.7415 | 10200
10210 | 48.2583 | 23.7407 | 10210
10220 | 48.47 | 23.7416 | 10220
10230 | 47.9995 | 23.7408 | 10230
10240 | 47.9243 | 23.7406 | 10240
10250 | 48.1124 | 23.7441 | 10250
10260 | 48.2818 | 23.7407 | 10260
10270 | 47.769 | 23.738 | 10270
10280 | 48.2818 | 23.7403 | 10280
10290 | 47.9384 | 23.7397 | 10290
10300 | 47.7361 | 23.7389 | 10300
10310 | 48.263 | 23.7367 | 10310
10320 | 48.0372 | 23.7478 | 10320
10330 | 47.9713 | 23.7387 | 10330
10340 | 48.4276 | 23.7381 | 10340
10350 | 48.3335 | 23.7412 | 10350
10360 | 47.8678 | 23.7378 | 10360
10370 | 48.4135 | 23.7412 | 10370
10380 | 48.1124 | 23.7388 | 10380
10390 | 48.0889 | 23.7419 | 10390
10400 | 48.3147 | 23.7441 | 10400
10410 | 47.5009 | 23.7379 | 10410
10420 | 48.2206 | 23.7368 | 10420
10430 | 48.4888 | 23.742 | 10430
10440 | 47.8866 | 23.7384 | 10440
10450 | 48.1689 | 23.7393 | 10450
10460 | 47.9572 | 23.7379 | 10460
10470 | 48.1689 | 23.7418 | 10470
10480 | 47.8866 | 23.7382 | 10480
10490 | 48.1124 | 23.7396 | 10490
10500 | 47.8584 | 23.7418 | 10500
10510 | 48.1454 | 23.7372 | 10510
10520 | 48.0654 | 23.746 | 10520
10530 | 48.2442 | 23.738 | 10530
10540 | 48.0184 | 23.7371 | 10540
10550 | 47.9243 | 23.7427 | 10550
10560 | 48.4323 | 23.7372 | 10560
10570 | 48.3524 | 23.7432 | 10570
10580 | 48.1689 | 23.7419 | 10580
10590 | 48.0513 | 23.7378 | 10590
10600 | 48.0889 | 23.741 | 10600
10610 | 48.1454 | 23.7385 | 10610
10620 | 47.9619 | 23.7349 | 10620
10630 | 48.2018 | 23.7412 | 10630
10640 | 48.1077 | 23.7374 | 10640
10650 | 47.6608 | 23.7477 | 10650
10660 | 47.769 | 23.7393 | 10660
10670 | 48.1877 | 23.7401 | 10670
10680 | 48.1689 | 23.7401 | 10680
10690 | 48.2206 | 23.7406 | 10690
10700 | 47.9807 | 23.7363 | 10700
10710 | 48.4653 | 23.7438 | 10710
10720 | 47.9995 | 23.7387 | 10720
10730 | 48.0136 | 23.7392 | 10730
10740 | 48.4888 | 23.7397 | 10740
10750 | 48.1266 | 23.7433 | 10750
10760 | 47.8866 | 23.7349 | 10760
10770 | 48.3947 | 23.7365 | 10770
10780 | 47.9572 | 23.7362 | 10780
10790 | 47.9007 | 23.7378 | 10790
10800 | 47.9054 | 23.7441 | 10800
10810 | 48.1124 | 23.7453 | 10810
10820 | 47.9384 | 23.7395 | 10820
10830 | 48.1454 | 23.7367 | 10830
10840 | 47.7925 | 23.7351 | 10840
10850 | 47.9054 | 23.7372 | 10850
10860 | 47.6373 | 23.7355 | 10860
10870 | 48.0936 | 23.7364 | 10870
10880 | 48.183 | 23.7378 | 10880
10890 | 48.056 | 23.744 | 10890
10900 | 48.1454 | 23.7466 | 10900
10910 | 48.2442 | 23.7372 | 10910
10920 | 48.47 | 23.7401 | 10920
10930 | 48.2065 | 23.7369 | 10930
10940 | 48.0701 | 23.7383 | 10940
10950 | 48.2253 | 23.7382 | 10950
10960 | 47.8631 | 23.7384 | 10960
10970 | 48.0748 | 23.7417 | 10970
10980 | 48.0372 | 23.7426 | 10980
10990 | 48.1454 | 23.7411 | 10990
11000 | 48.1642 | 23.7419 | 11000
11010 | 48.2818 | 23.7449 | 11010
11020 | 47.9243 | 23.7399 | 11020
11030 | 48.2395 | 23.7406 | 11030
11040 | 47.7173 | 23.7422 | 11040
11050 | 48.1077 | 23.739 | 11050
11060 | 47.8302 | 23.7391 | 11060
11070 | 48.263 | 23.7371 | 11070
11080 | 47.7173 | 23.7359 | 11080
11090 | 47.802 | 23.7452 | 11090
11100 | 47.9572 | 23.7452 | 11100
11110 | 47.7361 | 23.7411 | 11110
11120 | 48.056 | 23.74 | 11120
11130 | 47.7361 | 23.7429 | 11130
11140 | 48.2065 | 23.7432 | 11140
11150 | 47.8819 | 23.7388 | 11150
11160 | 48.9404 | 23.7375 | 11160
11170 | 47.9807 | 23.7474 | 11170
11180 | 47.8678 | 23.7353 | 11180
11190 | 48.1689 | 23.7446 | 11190
11200 | 47.7361 | 23.7384 | 11200
11210 | 48.4888 | 23.7425 | 11210
11220 | 48.0184 | 23.7378 | 11220
11230 | 48.2018 | 23.7405 | 11230
11240 | 47.8255 | 23.7361 | 11240
11250 | 48.1077 | 23.745 | 11250
11260 | 48.1454 | 23.7412 | 11260
11270 | 48.4888 | 23.7409 | 11270
11280 | 47.8067 | 23.7377 | 11280
11290 | 48.3759 | 23.7392 | 11290
11300 | 48.0748 | 23.7377 | 11300
11310 | 48.1266 | 23.7396 | 11310
11320 | 48.0513 | 23.7451 | 11320
11330 | 48.1642 | 23.7391 | 11330
11340 | 47.8302 | 23.7415 | 11340
11350 | 47.9431 | 23.7411 | 11350
11360 | 48.1077 | 23.737 | 11360
11370 | 47.976 | 23.7376 | 11370
11380 | 48.103 | 23.7365 | 11380
11390 | 48.0748 | 23.7382 | 11390
11400 | 48.103 | 23.7388 | 11400
11410 | 47.9619 | 23.7446 | 11410
11420 | 48.0513 | 23.7421 | 11420
11430 | 47.7314 | 23.7399 | 11430
11440 | 47.9619 | 23.7408 | 11440
11450 | 47.6561 | 23.7362 | 11450
11460 | 48.2442 | 23.7428 | 11460
11470 | 48.0889 | 23.7429 | 11470
11480 | 48.0701 | 23.7391 | 11480
11490 | 48.0889 | 23.7434 | 11490
11500 | 47.9995 | 23.735 | 11500
11510 | 48.0701 | 23.7367 | 11510
11520 | 47.9807 | 23.7401 | 11520
11530 | 48.0136 | 23.7401 | 11530
11540 | 48.2018 | 23.7382 | 11540
11550 | 48.5076 | 23.743 | 11550
11560 | 47.9149 | 23.7347 | 11560
11570 | 48.056 | 23.7354 | 11570
11580 | 48.1642 | 23.7426 | 11580
11590 | 48.2018 | 23.7371 | 11590
11600 | 47.9948 | 23.74 | 11600
11610 | 47.9713 | 23.7398 | 11610
11620 | 48.39 | 23.7375 | 11620
11630 | 48.4794 | 23.7393 | 11630
11640 | 47.9196 | 23.7403 | 11640
11650 | 48.0325 | 23.7354 | 11650
11660 | 48.2583 | 23.737 | 11660
11670 | 48.0748 | 23.74 | 11670
11680 | 47.6796 | 23.7369 | 11680
11690 | 48.1077 | 23.7387 | 11690
11700 | 47.8067 | 23.7334 | 11700
11710 | 48.5641 | 23.7397 | 11710
11720 | 47.769 | 23.7347 | 11720
11730 | 47.7361 | 23.7343 | 11730
11740 | 48.1266 | 23.7383 | 11740
11750 | 48.7099 | 23.7333 | 11750
11760 | 48.0748 | 23.7331 | 11760
11770 | 47.6796 | 23.7342 | 11770
11780 | 48.0889 | 23.7373 | 11780
11790 | 48.0513 | 23.7334 | 11790
11800 | 47.8631 | 23.7415 | 11800
11810 | 48.0325 | 23.7383 | 11810
11820 | 48.1642 | 23.7399 | 11820
11830 | 48.3524 | 23.7419 | 11830
11840 | 48.0513 | 23.7375 | 11840
11850 | 47.8866 | 23.7408 | 11850
11860 | 48.2018 | 23.7425 | 11860
11870 | 47.6185 | 23.7379 | 11870
11880 | 47.9007 | 23.7377 | 11880
11890 | 48.2018 | 23.7409 | 11890
11900 | 48.2959 | 23.7364 | 11900
11910 | 48.47 | 23.7377 | 11910
11920 | 47.6796 | 23.7401 | 11920
11930 | 48.2253 | 23.7405 | 11930
11940 | 48.0748 | 23.7323 | 11940
11950 | 47.4303 | 23.7316 | 11950
11960 | 47.7737 | 23.736 | 11960
11970 | 48.1454 | 23.7405 | 11970
11980 | 48.2253 | 23.7377 | 11980
11990 | 48.5029 | 23.7346 | 11990
12000 | 48.3147 | 23.7387 | 12000
12010 | 48.2018 | 23.7403 | 12010
12020 | 47.8067 | 23.7339 | 12020
12030 | 48.1454 | 23.7358 | 12030
12040 | 48.2818 | 23.7361 | 12040
12050 | 47.9431 | 23.7339 | 12050
12060 | 48.0889 | 23.7364 | 12060
12070 | 48.0513 | 23.7346 | 12070
12080 | 47.976 | 23.7378 | 12080
12090 | 47.8678 | 23.7385 | 12090
12100 | 47.6561 | 23.7362 | 12100
12110 | 48.2395 | 23.7376 | 12110
12120 | 48.0184 | 23.735 | 12120
12130 | 47.7549 | 23.7408 | 12130
12140 | 48.0936 | 23.7363 | 12140
12150 | 48.263 | 23.742 | 12150
12160 | 48.47 | 23.7344 | 12160
12170 | 48.1454 | 23.7337 | 12170
12180 | 47.6608 | 23.7387 | 12180
12190 | 48.4512 | 23.7365 | 12190
12200 | 47.7361 | 23.7334 | 12200
12210 | 47.8443 | 23.7334 | 12210
12220 | 48.0936 | 23.7326 | 12220
12230 | 48.4135 | 23.7378 | 12230
12240 | 48.2536 | 23.737 | 12240
12250 | 48.263 | 23.7327 | 12250
12260 | 47.9243 | 23.738 | 12260
12270 | 48.2395 | 23.7326 | 12270
12280 | 47.5056 | 23.7365 | 12280
12290 | 47.4726 | 23.7363 | 12290
12300 | 48.0889 | 23.7344 | 12300
12310 | 48.263 | 23.7364 | 12310
12320 | 48.2253 | 23.7408 | 12320
12330 | 47.8866 | 23.7369 | 12330
12340 | 48.4135 | 23.7375 | 12340
12350 | 47.802 | 23.7356 | 12350
12360 | 48.0748 | 23.7325 | 12360
12370 | 48.3853 | 23.7339 | 12370
12380 | 47.5856 | 23.7352 | 12380
12390 | 47.6749 | 23.7344 | 12390
12400 | 48.2253 | 23.7354 | 12400
12410 | 47.6561 | 23.7431 | 12410
12420 | 47.9007 | 23.7337 | 12420
12430 | 48.2583 | 23.7327 | 12430
12440 | 47.9243 | 23.7364 | 12440
12450 | 47.8866 | 23.7336 | 12450
12460 | 47.8302 | 23.733 | 12460
12470 | 48.1124 | 23.7406 | 12470
12480 | 48.1124 | 23.7322 | 12480
12490 | 48.2065 | 23.7374 | 12490
12500 | 47.8255 | 23.7351 | 12500
12510 | 48.2583 | 23.7337 | 12510
12520 | 48.3571 | 23.7333 | 12520
12530 | 47.9196 | 23.7375 | 12530
12540 | 47.8114 | 23.7353 | 12540
12550 | 48.1454 | 23.7359 | 12550
12560 | 48.2442 | 23.738 | 12560
12570 | 48.0513 | 23.7394 | 12570
12580 | 48.183 | 23.7326 | 12580
12590 | 47.8678 | 23.7344 | 12590
12600 | 48.0325 | 23.7352 | 12600
12610 | 47.8067 | 23.7356 | 12610
12620 | 48.2018 | 23.734 | 12620
12630 | 48.2771 | 23.737 | 12630
12640 | 47.8631 | 23.7366 | 12640
12650 | 47.8866 | 23.7385 | 12650
12660 | 47.9572 | 23.7321 | 12660
12670 | 48.1501 | 23.7335 | 12670
12680 | 47.5056 | 23.7377 | 12680
12690 | 48.2583 | 23.7366 | 12690
12700 | 47.9384 | 23.7384 | 12700
12710 | 48.1124 | 23.7323 | 12710
12720 | 48.1642 | 23.7363 | 12720
12730 | 48.1454 | 23.7352 | 12730
12740 | 48.263 | 23.7373 | 12740
12750 | 47.8443 | 23.7337 | 12750
12760 | 48.0936 | 23.7341 | 12760
12770 | 48.39 | 23.7373 | 12770
12780 | 47.9054 | 23.7377 | 12780
12790 | 48.0889 | 23.7362 | 12790
12800 | 47.6749 | 23.7344 | 12800
12810 | 48.39 | 23.7323 | 12810
12820 | 47.8114 | 23.7322 | 12820
12830 | 47.9572 | 23.7387 | 12830
12840 | 48.0513 | 23.7404 | 12840
12850 | 48.3524 | 23.7311 | 12850
12860 | 47.5761 | 23.7357 | 12860
12870 | 47.9948 | 23.7393 | 12870
12880 | 48.0701 | 23.7344 | 12880
12890 | 48.1077 | 23.7349 | 12890
12900 | 48.3524 | 23.7385 | 12900
12910 | 48.263 | 23.7351 | 12910
12920 | 48.4512 | 23.7397 | 12920
12930 | 48.0325 | 23.7393 | 12930
12940 | 47.8866 | 23.7385 | 12940
12950 | 48.2018 | 23.7384 | 12950
12960 | 48.0748 | 23.7344 | 12960
12970 | 48.0748 | 23.7338 | 12970
12980 | 48.056 | 23.7376 | 12980
12990 | 47.8866 | 23.7397 | 12990
13000 | 48.1689 | 23.7406 | 13000
13010 | 47.9243 | 23.7408 | 13010
13020 | 47.8631 | 23.7395 | 13020
13030 | 48.5358 | 23.7372 | 13030
13040 | 48.2442 | 23.7394 | 13040
13050 | 48.3194 | 23.7351 | 13050
13060 | 47.8302 | 23.7412 | 13060
13070 | 48.4512 | 23.7366 | 13070
13080 | 47.6373 | 23.7367 | 13080
13090 | 47.435 | 23.7346 | 13090
13100 | 48.0701 | 23.7367 | 13100
13110 | 48.2959 | 23.7316 | 13110
13120 | 48.263 | 23.7375 | 13120
13130 | 47.5856 | 23.7328 | 13130
13140 | 47.5667 | 23.7347 | 13140
13150 | 48.0513 | 23.7363 | 13150
13160 | 47.9196 | 23.7337 | 13160
13170 | 47.849 | 23.7358 | 13170
13180 | 48.3524 | 23.7425 | 13180
13190 | 48.1454 | 23.7361 | 13190
13200 | 47.8114 | 23.7356 | 13200
13210 | 47.9995 | 23.7356 | 13210
13220 | 47.9384 | 23.7348 | 13220
13230 | 48.056 | 23.7334 | 13230
13240 | 48.0748 | 23.7386 | 13240
13250 | 47.9572 | 23.7342 | 13250
13260 | 48.2253 | 23.7326 | 13260
13270 | 48.0748 | 23.7375 | 13270
13280 | 47.9572 | 23.7377 | 13280
13290 | 47.9431 | 23.7318 | 13290
13300 | 48.3006 | 23.7352 | 13300
13310 | 48.3947 | 23.734 | 13310
13320 | 48.0513 | 23.7401 | 13320
13330 | 47.8678 | 23.733 | 13330
13340 | 48.1689 | 23.7333 | 13340
13350 | 48.183 | 23.7393 | 13350
13360 | 48.3759 | 23.7334 | 13360
13370 | 48.0184 | 23.7401 | 13370
13380 | 48.3006 | 23.7406 | 13380
13390 | 47.7737 | 23.7347 | 13390
13400 | 48.0184 | 23.734 | 13400
13410 | 47.9572 | 23.7392 | 13410
13420 | 48.1642 | 23.7331 | 13420
13430 | 48.39 | 23.7333 | 13430
13440 | 48.0184 | 23.7397 | 13440
13450 | 48.1971 | 23.7332 | 13450
13460 | 48.2018 | 23.7327 | 13460
13470 | 48.0748 | 23.7328 | 13470
13480 | 48.0748 | 23.7305 | 13480
13490 | 47.7173 | 23.7375 | 13490
13500 | 48.1642 | 23.7342 | 13500
13510 | 48.2959 | 23.7391 | 13510
13520 | 47.7314 | 23.734 | 13520
13530 | 48.1454 | 23.7379 | 13530
13540 | 47.9149 | 23.737 | 13540
13550 | 48.2771 | 23.7327 | 13550
13560 | 47.7737 | 23.7294 | 13560
13570 | 48.1501 | 23.7325 | 13570
13580 | 48.1124 | 23.7356 | 13580
13590 | 48.0136 | 23.7385 | 13590
13600 | 47.6185 | 23.7333 | 13600
13610 | 48.056 | 23.7326 | 13610
13620 | 47.9807 | 23.7313 | 13620
13630 | 48.1501 | 23.7294 | 13630
13640 | 48.1689 | 23.7324 | 13640
13650 | 48.056 | 23.7316 | 13650
13660 | 48.1077 | 23.7381 | 13660
13670 | 47.9807 | 23.7344 | 13670
13680 | 47.6185 | 23.7384 | 13680
13690 | 48.1501 | 23.739 | 13690
13700 | 48.2959 | 23.7332 | 13700
13710 | 48.1642 | 23.7321 | 13710
13720 | 47.9337 | 23.7325 | 13720
13730 | 48.597 | 23.7339 | 13730
13740 | 48.1877 | 23.7392 | 13740
13750 | 48.0513 | 23.7346 | 13750
13760 | 47.9619 | 23.7331 | 13760
13770 | 48.2818 | 23.7369 | 13770
13780 | 48.263 | 23.7344 | 13780
13790 | 48.0748 | 23.7377 | 13790
13800 | 47.7878 | 23.7317 | 13800
13810 | 47.7314 | 23.7386 | 13810
13820 | 48.3194 | 23.7326 | 13820
13830 | 47.9995 | 23.7361 | 13830
13840 | 47.9948 | 23.7319 | 13840
13850 | 47.7314 | 23.7305 | 13850
13860 | 47.7126 | 23.7333 | 13860
13870 | 48.2206 | 23.7369 | 13870
13880 | 47.9243 | 23.736 | 13880
13890 | 48.4088 | 23.736 | 13890
13900 | 47.8067 | 23.7311 | 13900
13910 | 48.1877 | 23.7316 | 13910
13920 | 48.3382 | 23.7341 | 13920
13930 | 48.1313 | 23.7375 | 13930
13940 | 47.802 | 23.7309 | 13940
13950 | 47.9948 | 23.7306 | 13950
13960 | 48.0513 | 23.73 | 13960
13970 | 47.7925 | 23.7379 | 13970
13980 | 48.263 | 23.7325 | 13980
13990 | 48.0513 | 23.7342 | 13990
14000 | 48.0513 | 23.732 | 14000
14010 | 48.1124 | 23.7354 | 14010
14020 | 48.3147 | 23.7358 | 14020
14030 | 48.4464 | 23.7374 | 14030
14040 | 48.1218 | 23.7293 | 14040
14050 | 48.3853 | 23.732 | 14050
14060 | 48.4135 | 23.7296 | 14060
14070 | 47.9196 | 23.7294 | 14070
14080 | 48.5594 | 23.7346 | 14080
14090 | 48.3194 | 23.7364 | 14090
14100 | 48.5076 | 23.7309 | 14100
14110 | 48.0748 | 23.7321 | 14110
14120 | 48.0513 | 23.7292 | 14120
14130 | 47.9948 | 23.73 | 14130
14140 | 48.5029 | 23.7356 | 14140
14150 | 47.9243 | 23.7328 | 14150
14160 | 48.1266 | 23.7337 | 14160
14170 | 48.0136 | 23.7386 | 14170
14180 | 47.9384 | 23.7315 | 14180
14190 | 48.1313 | 23.7338 | 14190
14200 | 47.9007 | 23.7333 | 14200
14210 | 48.2395 | 23.7367 | 14210
14220 | 48.1077 | 23.7308 | 14220
14230 | 48.263 | 23.7337 | 14230
14240 | 48.4276 | 23.7301 | 14240
14250 | 48.0184 | 23.7362 | 14250
14260 | 48.1313 | 23.7302 | 14260
14270 | 47.9948 | 23.7359 | 14270
14280 | 47.849 | 23.7323 | 14280
14290 | 48.263 | 23.7387 | 14290
14300 | 48.183 | 23.733 | 14300
14310 | 47.8678 | 23.7398 | 14310
14320 | 48.1689 | 23.7366 | 14320
14330 | 47.7173 | 23.7379 | 14330
14340 | 48.0184 | 23.7368 | 14340
14350 | 47.6561 | 23.736 | 14350
14360 | 48.3335 | 23.7384 | 14360
14370 | 48.2065 | 23.7365 | 14370
14380 | 48.5405 | 23.7354 | 14380
14390 | 48.0325 | 23.7329 | 14390
14400 | 47.9948 | 23.734 | 14400
14410 | 47.9807 | 23.7326 | 14410
14420 | 48.1689 | 23.7376 | 14420
14430 | 47.8631 | 23.7339 | 14430
14440 | 48.4135 | 23.7352 | 14440
14450 | 47.9525 | 23.7311 | 14450
14460 | 48.2018 | 23.7338 | 14460
14470 | 48.0184 | 23.7314 | 14470
14480 | 48.4464 | 23.736 | 14480
14490 | 48.0325 | 23.7311 | 14490
14500 | 48.3712 | 23.7348 | 14500
14510 | 48.3947 | 23.7322 | 14510
14520 | 47.8819 | 23.7307 | 14520
14530 | 48.183 | 23.7345 | 14530
14540 | 48.47 | 23.7347 | 14540
14550 | 47.7925 | 23.7344 | 14550
14560 | 47.9572 | 23.732 | 14560
14570 | 47.6044 | 23.7342 | 14570
14580 | 48.1266 | 23.737 | 14580
14590 | 48.103 | 23.7364 | 14590
14600 | 47.9007 | 23.7314 | 14600
14610 | 48.056 | 23.73 | 14610
14620 | 47.7925 | 23.734 | 14620
14630 | 47.7925 | 23.7318 | 14630
14640 | 47.8067 | 23.7308 | 14640
14650 | 48.1689 | 23.7322 | 14650
14660 | 48.0136 | 23.7299 | 14660
14670 | 48.263 | 23.7299 | 14670
14680 | 48.2018 | 23.7358 | 14680
14690 | 48.5594 | 23.7349 | 14690
14700 | 48.0513 | 23.7316 | 14700
14710 | 48.1454 | 23.7315 | 14710
14720 | 48.1266 | 23.7366 | 14720
14730 | 48.4512 | 23.7357 | 14730
14740 | 48.1218 | 23.7317 | 14740
14750 | 47.9196 | 23.7316 | 14750
14760 | 47.976 | 23.7335 | 14760
14770 | 48.4888 | 23.7349 | 14770
14780 | 48.2018 | 23.7382 | 14780
14790 | 48.1124 | 23.7332 | 14790
14800 | 47.9384 | 23.732 | 14800
14810 | 47.9007 | 23.7344 | 14810
14820 | 48.0513 | 23.7314 | 14820
14830 | 48.1642 | 23.731 | 14830
14840 | 48.2818 | 23.7354 | 14840
14850 | 48.0889 | 23.7311 | 14850
14860 | 47.9619 | 23.7345 | 14860
14870 | 47.9384 | 23.728 | 14870
14880 | 47.8866 | 23.7345 | 14880
14890 | 48.0325 | 23.7337 | 14890
14900 | 47.8866 | 23.7335 | 14900
14910 | 48.0136 | 23.7305 | 14910
14920 | 47.9196 | 23.7321 | 14920
14930 | 48.0184 | 23.7349 | 14930
14940 | 48.0184 | 23.7321 | 14940
14950 | 48.1313 | 23.7299 | 14950
14960 | 47.8866 | 23.7316 | 14960
14970 | 48.5217 | 23.7321 | 14970
14980 | 47.6796 | 23.7266 | 14980
14990 | 48.4841 | 23.7346 | 14990
15000 | 48.2583 | 23.7311 | 15000
15010 | 47.976 | 23.7319 | 15010
15020 | 48.4464 | 23.7354 | 15020
15030 | 47.8866 | 23.7297 | 15030
15040 | 48.0136 | 23.7301 | 15040
15050 | 48.3524 | 23.7294 | 15050
15060 | 48.0513 | 23.7312 | 15060
15070 | 47.849 | 23.7291 | 15070
15080 | 48.1077 | 23.7322 | 15080
15090 | 47.5244 | 23.7293 | 15090
15100 | 47.9384 | 23.7419 | 15100
15110 | 47.9572 | 23.7382 | 15110
15120 | 47.5573 | 23.7343 | 15120
15130 | 47.8866 | 23.7297 | 15130
15140 | 47.7643 | 23.7326 | 15140
15150 | 48.0936 | 23.7285 | 15150
15160 | 48.1971 | 23.7311 | 15160
15170 | 48.1077 | 23.7325 | 15170
15180 | 48.1124 | 23.7297 | 15180
15190 | 48.4135 | 23.7271 | 15190
15200 | 48.2442 | 23.7351 | 15200
15210 | 48.1454 | 23.7283 | 15210
15220 | 48.1266 | 23.7273 | 15220
15230 | 48.4135 | 23.7301 | 15230
15240 | 47.8396 | 23.7343 | 15240
15250 | 48.0278 | 23.7292 | 15250
15260 | 47.8255 | 23.7285 | 15260
15270 | 48.0513 | 23.7356 | 15270
15280 | 48.2395 | 23.7342 | 15280
15290 | 48.0136 | 23.7325 | 15290
15300 | 48.4135 | 23.7339 | 15300
15310 | 48.4088 | 23.7306 | 15310
15320 | 47.9995 | 23.7276 | 15320
15330 | 48.1124 | 23.7391 | 15330
15340 | 47.6561 | 23.7301 | 15340
15350 | 48.47 | 23.7367 | 15350
15360 | 48.1407 | 23.7314 | 15360
15370 | 48.1454 | 23.7299 | 15370
15380 | 47.976 | 23.7264 | 15380
15390 | 48.1077 | 23.7324 | 15390
15400 | 47.8819 | 23.7315 | 15400
15410 | 48.1266 | 23.737 | 15410
15420 | 48.2583 | 23.7293 | 15420
15430 | 47.7173 | 23.7308 | 15430
15440 | 48.1454 | 23.7288 | 15440
15450 | 48.3712 | 23.7285 | 15450
15460 | 47.8631 | 23.7316 | 15460
15470 | 47.642 | 23.7353 | 15470
15480 | 47.9054 | 23.7304 | 15480
15490 | 48.2771 | 23.729 | 15490
15500 | 47.6938 | 23.7316 | 15500
15510 | 48.0748 | 23.7282 | 15510
15520 | 48.1407 | 23.7302 | 15520
15530 | 48.5264 | 23.729 | 15530
15540 | 48.3335 | 23.732 | 15540
15550 | 48.2065 | 23.728 | 15550
15560 | 48.1313 | 23.729 | 15560
15570 | 47.976 | 23.7268 | 15570
15580 | 48.2959 | 23.7288 | 15580
15590 | 47.8114 | 23.7273 | 15590
15600 | 47.8678 | 23.7303 | 15600
15610 | 47.7455 | 23.7302 | 15610
15620 | 47.8208 | 23.7295 | 15620
15630 | 47.8067 | 23.7338 | 15630
15640 | 48.0889 | 23.7305 | 15640
15650 | 47.6608 | 23.738 | 15650
15660 | 48.1124 | 23.7289 | 15660
15670 | 47.7455 | 23.7283 | 15670
15680 | 48.5076 | 23.7314 | 15680
15690 | 48.0748 | 23.7318 | 15690
15700 | 48.4323 | 23.7331 | 15700
15710 | 48.3006 | 23.7282 | 15710
15720 | 48.0136 | 23.7308 | 15720
15730 | 47.8819 | 23.7327 | 15730
15740 | 47.8443 | 23.7319 | 15740
15750 | 47.9807 | 23.734 | 15750
15760 | 48.4512 | 23.7302 | 15760
15770 | 48.2442 | 23.7354 | 15770
15780 | 48.2959 | 23.7328 | 15780
15790 | 48.2583 | 23.7387 | 15790
15800 | 48.0136 | 23.73 | 15800
15810 | 48.2442 | 23.7326 | 15810
15820 | 48.1313 | 23.7336 | 15820
15830 | 48.1454 | 23.7292 | 15830
15840 | 48.3947 | 23.7309 | 15840
15850 | 48.1077 | 23.7291 | 15850
15860 | 48.263 | 23.7286 | 15860
15870 | 48.3006 | 23.7279 | 15870
15880 | 48.1501 | 23.7291 | 15880
15890 | 48.2912 | 23.731 | 15890
15900 | 47.9807 | 23.7261 | 15900
15910 | 47.9807 | 23.7343 | 15910
15920 | 48.3571 | 23.7319 | 15920
15930 | 47.8631 | 23.7318 | 15930
15940 | 47.7549 | 23.7277 | 15940
15950 | 48.0372 | 23.7328 | 15950
15960 | 47.8866 | 23.7297 | 15960
15970 | 48.2253 | 23.7286 | 15970
15980 | 47.8866 | 23.7256 | 15980
15990 | 48.39 | 23.7277 | 15990
16000 | 47.9619 | 23.7309 | 16000
16010 | 48.0136 | 23.7302 | 16010
16020 | 48.0748 | 23.7279 | 16020
16030 | 47.7079 | 23.7314 | 16030
16040 | 48.3759 | 23.7263 | 16040
16050 | 47.9431 | 23.7286 | 16050
16060 | 48.4323 | 23.7254 | 16060
16070 | 48.0136 | 23.7304 | 16070
16080 | 47.7173 | 23.7349 | 16080
16090 | 48.1124 | 23.7345 | 16090
16100 | 47.8631 | 23.7286 | 16100
16110 | 48.2253 | 23.73 | 16110
16120 | 47.642 | 23.7313 | 16120
16130 | 48.4323 | 23.7285 | 16130
16140 | 47.5808 | 23.729 | 16140
16150 | 47.6985 | 23.7299 | 16150
16160 | 48.3571 | 23.7274 | 16160
16170 | 47.8678 | 23.7261 | 16170
16180 | 47.9572 | 23.7306 | 16180
16190 | 47.976 | 23.7275 | 16190
16200 | 48.2583 | 23.7295 | 16200
16210 | 48.0184 | 23.7305 | 16210
16220 | 48.1501 | 23.7305 | 16220
16230 | 48.0748 | 23.7321 | 16230
16240 | 48.2442 | 23.727 | 16240
16250 | 48.0701 | 23.7302 | 16250
16260 | 47.9807 | 23.7304 | 16260
16270 | 47.9054 | 23.7289 | 16270
16280 | 47.7173 | 23.7291 | 16280
16290 | 48.1642 | 23.7343 | 16290
16300 | 48.3335 | 23.7328 | 16300
16310 | 48.4653 | 23.7313 | 16310
16320 | 47.7643 | 23.7311 | 16320
16330 | 48.2818 | 23.7281 | 16330
16340 | 48.1407 | 23.7273 | 16340
16350 | 48.3571 | 23.7285 | 16350
16360 | 47.8678 | 23.7291 | 16360
16370 | 48.5076 | 23.729 | 16370
16380 | 47.9948 | 23.7345 | 16380
16390 | 47.7643 | 23.7265 | 16390
16400 | 48.3006 | 23.7317 | 16400
16410 | 48.0372 | 23.7271 | 16410
16420 | 48.1689 | 23.7255 | 16420
16430 | 47.7361 | 23.7307 | 16430
16440 | 47.9995 | 23.7351 | 16440
16450 | 47.9196 | 23.7313 | 16450
16460 | 48.056 | 23.7258 | 16460
16470 | 48.47 | 23.7316 | 16470
16480 | 47.9619 | 23.7288 | 16480
16490 | 48.2818 | 23.7305 | 16490
16500 | 47.8302 | 23.7317 | 16500
16510 | 48.0701 | 23.733 | 16510
16520 | 47.8866 | 23.7272 | 16520
16530 | 47.9807 | 23.7305 | 16530
16540 | 48.2395 | 23.7262 | 16540
16550 | 48.5641 | 23.729 | 16550
16560 | 47.8067 | 23.7302 | 16560
16570 | 47.8631 | 23.7299 | 16570
16580 | 47.9995 | 23.733 | 16580
16590 | 47.8678 | 23.7278 | 16590
16600 | 48.0184 | 23.7311 | 16600
16610 | 47.6796 | 23.7323 | 16610
16620 | 48.1266 | 23.7268 | 16620
16630 | 48.1077 | 23.7294 | 16630
16640 | 47.7361 | 23.7289 | 16640
16650 | 48.1454 | 23.7286 | 16650
16660 | 48.1642 | 23.7293 | 16660
16670 | 48.0184 | 23.7321 | 16670
16680 | 47.8584 | 23.7299 | 16680
16690 | 48.2253 | 23.729 | 16690
16700 | 48.0184 | 23.7266 | 16700
16710 | 47.8631 | 23.7261 | 16710
16720 | 48.1077 | 23.7252 | 16720
16730 | 47.7737 | 23.7252 | 16730
16740 | 48.2065 | 23.7264 | 16740
16750 | 48.1689 | 23.7331 | 16750
16760 | 48.1783 | 23.73 | 16760
16770 | 48.0513 | 23.7249 | 16770
16780 | 47.8302 | 23.7269 | 16780
16790 | 48.0748 | 23.7277 | 16790
16800 | 47.8819 | 23.7288 | 16800
16810 | 48.1077 | 23.7291 | 16810
16820 | 47.9431 | 23.7303 | 16820
16830 | 47.7737 | 23.7326 | 16830
16840 | 48.0748 | 23.7298 | 16840
16850 | 47.8631 | 23.7258 | 16850
16860 | 47.9054 | 23.7282 | 16860
16870 | 48.1877 | 23.7265 | 16870
16880 | 48.0936 | 23.7292 | 16880
16890 | 48.2018 | 23.7276 | 16890
16900 | 48.263 | 23.728 | 16900
16910 | 47.9196 | 23.7247 | 16910
16920 | 47.976 | 23.7253 | 16920
16930 | 48.0748 | 23.7311 | 16930
16940 | 48.1689 | 23.7302 | 16940
16950 | 48.056 | 23.73 | 16950
16960 | 47.8631 | 23.7332 | 16960
16970 | 48.0513 | 23.7308 | 16970
16980 | 48.0701 | 23.7276 | 16980
16990 | 47.9572 | 23.7288 | 16990
17000 | 48.0136 | 23.733 | 17000
17010 | 48.5405 | 23.7322 | 17010
17020 | 48.056 | 23.7275 | 17020
17030 | 48.4088 | 23.7283 | 17030
17040 | 48.1266 | 23.7281 | 17040
17050 | 48.0701 | 23.7326 | 17050
17060 | 48.1595 | 23.7276 | 17060
17070 | 48.1454 | 23.7285 | 17070
17080 | 48.1877 | 23.7281 | 17080
17090 | 48.1313 | 23.7307 | 17090
17100 | 47.9196 | 23.7308 | 17100
17110 | 48.2771 | 23.7284 | 17110
17120 | 47.8678 | 23.7366 | 17120
17130 | 48.0184 | 23.7284 | 17130
17140 | 47.8584 | 23.7264 | 17140
17150 | 48.1077 | 23.7274 | 17150
17160 | 48.1642 | 23.7264 | 17160
17170 | 47.9243 | 23.7304 | 17170
17180 | 47.642 | 23.73 | 17180
17190 | 48.0889 | 23.7311 | 17190
17200 | 47.9572 | 23.7309 | 17200
17210 | 48.0136 | 23.7299 | 17210
17220 | 48.3712 | 23.7282 | 17220
17230 | 47.976 | 23.7263 | 17230
17240 | 48.1501 | 23.7284 | 17240
17250 | 47.8866 | 23.7245 | 17250
17260 | 47.849 | 23.7327 | 17260
17270 | 47.849 | 23.7338 | 17270
17280 | 47.9243 | 23.7249 | 17280
17290 | 47.9431 | 23.725 | 17290
17300 | 48.1266 | 23.7268 | 17300
17310 | 48.4135 | 23.7285 | 17310
17320 | 48.1642 | 23.7259 | 17320
17330 | 48.1266 | 23.7241 | 17330
17340 | 48.1501 | 23.7243 | 17340
17350 | 48.5029 | 23.7252 | 17350
17360 | 48.2771 | 23.7281 | 17360
17370 | 48.1642 | 23.7298 | 17370
17380 | 48.3524 | 23.7272 | 17380
17390 | 48.3524 | 23.7253 | 17390
17400 | 47.6985 | 23.7304 | 17400
17410 | 48.597 | 23.7264 | 17410
17420 | 47.7314 | 23.7254 | 17420
17430 | 48.263 | 23.7281 | 17430
17440 | 48.1077 | 23.7264 | 17440
17450 | 48.183 | 23.7274 | 17450
17460 | 48.1313 | 23.7296 | 17460
17470 | 48.4135 | 23.7281 | 17470
17480 | 48.0889 | 23.7303 | 17480
17490 | 48.0889 | 23.727 | 17490
17500 | 47.976 | 23.7264 | 17500
17510 | 48.2959 | 23.7268 | 17510
17520 | 48.2065 | 23.7299 | 17520
17530 | 47.9995 | 23.7284 | 17530
17540 | 48.0184 | 23.7262 | 17540
17550 | 47.7549 | 23.7272 | 17550
17560 | 48.6158 | 23.7296 | 17560
17570 | 48.4135 | 23.729 | 17570
17580 | 48.2771 | 23.7258 | 17580
17590 | 48.0325 | 23.7271 | 17590
17600 | 48.0513 | 23.73 | 17600
17610 | 47.9243 | 23.7274 | 17610
17620 | 48.2442 | 23.7271 | 17620
17630 | 48.2206 | 23.7267 | 17630
17640 | 47.8067 | 23.726 | 17640
17650 | 48.1454 | 23.7298 | 17650
17660 | 48.056 | 23.7265 | 17660
17670 | 47.7878 | 23.7247 | 17670
17680 | 48.263 | 23.7287 | 17680
17690 | 48.263 | 23.726 | 17690
17700 | 48.4135 | 23.7314 | 17700
17710 | 48.0513 | 23.728 | 17710
17720 | 47.9807 | 23.7235 | 17720
17730 | 47.769 | 23.7268 | 17730
17740 | 48.2395 | 23.729 | 17740
17750 | 48.3571 | 23.7279 | 17750
17760 | 48.3006 | 23.724 | 17760
17770 | 47.9572 | 23.7264 | 17770
17780 | 47.9948 | 23.729 | 17780
17790 | 48.0136 | 23.7257 | 17790
17800 | 47.642 | 23.7257 | 17800
17810 | 47.9007 | 23.7285 | 17810
17820 | 48.47 | 23.7234 | 17820
17830 | 48.5029 | 23.7227 | 17830
17840 | 48.4276 | 23.7255 | 17840
17850 | 48.1266 | 23.7276 | 17850
17860 | 47.7173 | 23.7264 | 17860
17870 | 48.1642 | 23.7289 | 17870
17880 | 48.1266 | 23.7333 | 17880
17890 | 47.9619 | 23.7254 | 17890
17900 | 48.4135 | 23.7257 | 17900
17910 | 47.5479 | 23.7253 | 17910
17920 | 48.6205 | 23.7278 | 17920
17930 | 48.0325 | 23.7342 | 17930
17940 | 48.0136 | 23.7255 | 17940
17950 | 47.7079 | 23.7257 | 17950
17960 | 48.2018 | 23.7248 | 17960
17970 | 48.3194 | 23.7242 | 17970
17980 | 47.9619 | 23.7314 | 17980
17990 | 48.3288 | 23.7248 | 17990
18000 | 47.9995 | 23.7234 | 18000
18010 | 48.0136 | 23.7276 | 18010
18020 | 47.6796 | 23.7235 | 18020
18030 | 48.2253 | 23.7311 | 18030
18040 | 47.849 | 23.7247 | 18040
18050 | 47.4538 | 23.7246 | 18050
18060 | 48.1595 | 23.7284 | 18060
18070 | 47.6985 | 23.7282 | 18070
18080 | 48.3947 | 23.7269 | 18080
18090 | 48.263 | 23.7263 | 18090
18100 | 47.9384 | 23.727 | 18100
18110 | 47.9243 | 23.7276 | 18110
18120 | 48.4464 | 23.7213 | 18120
18130 | 48.2818 | 23.7261 | 18130
18140 | 47.9619 | 23.723 | 18140
18150 | 48.0513 | 23.7295 | 18150
18160 | 47.9619 | 23.7292 | 18160
18170 | 48.0701 | 23.7228 | 18170
18180 | 47.9948 | 23.7309 | 18180
18190 | 47.9196 | 23.7254 | 18190
18200 | 47.9572 | 23.7307 | 18200
18210 | 48.3477 | 23.7256 | 18210
18220 | 47.9007 | 23.7264 | 18220
18230 | 48.0889 | 23.7236 | 18230
18240 | 47.9007 | 23.7264 | 18240
18250 | 47.7737 | 23.7269 | 18250
18260 | 48.1501 | 23.7273 | 18260
18270 | 48.1642 | 23.7306 | 18270
18280 | 47.9619 | 23.7274 | 18280
18290 | 47.8255 | 23.7243 | 18290
18300 | 48.1313 | 23.7262 | 18300
18310 | 48.1642 | 23.7273 | 18310
18320 | 48.1124 | 23.7276 | 18320
18330 | 47.6796 | 23.7305 | 18330
18340 | 48.0372 | 23.726 | 18340
18350 | 47.9196 | 23.7328 | 18350
18360 | 48.1501 | 23.7263 | 18360
18370 | 48.0136 | 23.7266 | 18370
18380 | 48.2253 | 23.7287 | 18380
18390 | 47.9525 | 23.726 | 18390
18400 | 47.9948 | 23.7222 | 18400
18410 | 47.8631 | 23.7265 | 18410
18420 | 48.103 | 23.7264 | 18420
18430 | 48.3006 | 23.7285 | 18430
18440 | 47.8631 | 23.7233 | 18440
18450 | 48.1689 | 23.725 | 18450
18460 | 48.2442 | 23.7208 | 18460
18470 | 48.2206 | 23.7238 | 18470
18480 | 47.5667 | 23.7284 | 18480
18490 | 48.2395 | 23.7292 | 18490
18500 | 47.9901 | 23.7263 | 18500
18510 | 47.8819 | 23.7224 | 18510
18520 | 48.0748 | 23.7279 | 18520
18530 | 48.0372 | 23.7304 | 18530
18540 | 48.1642 | 23.7308 | 18540
18550 | 48.2959 | 23.7257 | 18550
18560 | 48.183 | 23.7228 | 18560
18570 | 48.3712 | 23.725 | 18570
18580 | 48.6158 | 23.7247 | 18580
18590 | 48.1595 | 23.7326 | 18590
18600 | 47.9572 | 23.7212 | 18600
18610 | 47.8819 | 23.7312 | 18610
18620 | 47.8302 | 23.7277 | 18620
18630 | 47.9054 | 23.729 | 18630
18640 | 48.3947 | 23.7251 | 18640
18650 | 48.0748 | 23.7306 | 18650
18660 | 48.1266 | 23.7246 | 18660
18670 | 47.8255 | 23.7272 | 18670
18680 | 47.9054 | 23.725 | 18680
18690 | 48.0513 | 23.7258 | 18690
18700 | 48.5029 | 23.7242 | 18700
18710 | 48.2253 | 23.7285 | 18710
18720 | 48.1313 | 23.7303 | 18720
18730 | 47.9243 | 23.727 | 18730
18740 | 47.7737 | 23.7255 | 18740
18750 | 47.9384 | 23.7296 | 18750
18760 | 48.0184 | 23.7268 | 18760
18770 | 48.2959 | 23.7304 | 18770
18780 | 47.8866 | 23.7234 | 18780
18790 | 48.183 | 23.7274 | 18790
18800 | 47.9243 | 23.7257 | 18800
18810 | 48.0748 | 23.7219 | 18810
Note on G4: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This is not used in the calculation, but is useful for identifying which load step this represents.
Note on B6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 27.819 kPa is the average stress applied to the sample for the previous loading step and is presented here to simplify the discussion. As this loads transitions to 48.1 kPa for this step, the load frame recorded a value at 29.577 (Cell B7).
Note on G7: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The first step is to select the time when the load step starts. To graph, select the row so that t=0 is the point just prior to the load increasing
Note on H7: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The row for d1 and t1 should be chosen to be before the onset of the more rapid stage of primary consolidation
Note on I7: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This is generally the last point in the consolidation, when consolidation is 100% complete.
Note on G10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Viscosity taken from the NISTbook value for 0.1 MPa (atmospheric pressure, since the oedometer is open to the atmosphere) and 22 C, the typical laboratory room temperature.
Note on J12: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This is taken as the average load applied during the previous load step.
Note on K12: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This is the average load for this loading step, taken from the final 1000 points in this step.
Note on P35: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This example is for the Frame 1, 0.05 MPa effective stress step shown in the “Permeability” tab.